## Sheet1
Utica 2.4 GHz Power Plus BTE 13 Metrics Report Grand Avenue Document ID: GQ-RPT-01045 Revision: 2.0
Description of data in this document: Function: This column designates which functional team is responsible for verifying the corresponding Systems requirement. Requirement Result: This column indicates the cumulative (across final test runs) verification result for the corresponding requirement. This is the cumulative result of the corresponding Test Resullt (Column H) data. Requirement: This column lists the requirement title as shown in Polarion and provides a hyperlink to the requirement in Polarion. Test Case: This column indicates which test case(s) contain steps that verify the requirement. Please note that Polarion includes all executed tests that reference the relevant requirement. Please also note that Polarion only includes verification results in this report from the most recently executed Test Run(s). For these reasons, some results included in this report may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Run: This column lists the test run(s) that contained the test case(s) (Column D) that verify the requirement. Please note that Polarion only includes verification results in this report from the most recently executed Test Run(s), which may or may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Result Status: This column indicates the execution status for each instance of the requirement in all executed test cases. Test Step: This column indicates the location of every instance of the requirement in the corresponding Test Case (Column D). Zero-step test cases (test cases that directly pass/fail an individual requirement) will indicate N/A in this column. Test Result: This column indicates the test result for every instance of the requirement in the corresponding Test Case (Column D). This column is the input for the Requirement Result (Column B) column.
Function | Requirement Result | Requirement | Test Case | Test Run | Test Result Status | Test Step | Test Result
Wireless | passed | MASTER-20000 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20001 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-48455 - The hearing aid shall be capable of communicating with compatible Starkey access... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20002 - With no line-of-sight from the BLE Central to the user's head due to body-blocki... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20011 - The hearing aid shall receive commands from the Remote Control at a distance of... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-64822 - For RIC and BTE styles, the hearing aid shall communicate E2E via the 2.4 GHz ra... | MASTER-62371 - E2E Hardware Performance - Standard Hearing Aids | Hearing Aid RF HW - P BTE | Executed | 3 | passed
Wireless | passed | MASTER-20027 - The hearing aid shall achieve a link availability of greater than or equal to 95... | MASTER-51619 - Hearing Aid Communication Reliability | Hearing Aid RF HW - P BTE | Executed | 2 | passed
Wireless | passed | MASTER-20049 - The hearing aid shall communicate wirelessly using Bluetooth Low Energy as defin... | MASTER-49133 - Bluetooth Low Energy Core Compliance | Hearing Aid RF Hardware - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20143 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20144 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 5 | passed
Wireless | passed | MASTER-20145 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 6 | passed
Wireless | passed | MASTER-20147 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 7 | passed
Wireless | passed | MASTER-20148 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 9 | passed
Wireless | passed | MASTER-20152 - The hearing aid shall comply with ERC/REC 70-03, Annex 3, frequency band A. | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 5 | passed
Wireless | passed | MASTER-20153 - The hearing aid shall comply with Commission Decision 2009/381/EC for operation... | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 6 | passed
Wireless | passed | MASTER-20154 - The hearing aid shall comply with ETSI EN 300 328 for operation within the band... | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20158 - The hearing aid shall comply with the requirements of RSS-Gen that are applicabl... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20159 - The hearing aid shall comply with RSS-247 for operation within the band of 2400-... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 6 | passed
Wireless | passed | MASTER-20160 - The hearing aid shall comply with RSS-102 for operation within the band of 2400-... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 9 | passed
Wireless | passed | MASTER-20161 - The hearing aid shall comply with RSP-100. | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 8 | passed
Wireless | passed | MASTER-20164 - The hearing aid shall comply with the performance and methods of measurement of... | MASTER-48794 - Australia Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20167 - The hearing aid shall comply with KN 301 489-1. Note: Testing is performed using... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 6 | passed
Wireless | passed | MASTER-42289 - The hearing aid shall comply with KN 301 489-17. Note: Testing is performed usin... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 5 | passed
Wireless | passed | MASTER-42290 - The hearing aid shall comply with Korean standard KN 300 328 for operation withi... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20170 - The hearing aid shall comply with China Ministry of Industry and Information Tec... | MASTER-48791 - China Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20172 - The hearing aid shall comply with Japan Association of Radio Industries and Busi... | MASTER-48790 - Japan Wireless Regulatory Conformance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20174 - The hearing aid shall be qualified to the standards defined by the Bluetooth Spe... | MASTER-49133 - Bluetooth Low Energy Core Compliance | Hearing Aid RF Hardware - CIC | Executed | 4 | passed
Firmware | passed | MASTER-19862 - When the Directionality Mode is Adaptive Directionality, the following audio sti... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19863 - When the Directionality Mode is Adaptive Directionality, the following audio sti... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-20021 - Latency for wireless data and command communication between E2E Partners shall n... | MASTER-54929 - E2E UI Command Latency | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Systems | passed | MASTER-17316 - The hearing aid shall be capable of electrical operation using a battery from al... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 21 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 74 | passed
Systems | passed | UTICA-171 - The hearing aid shall provide a battery life of greater than or equal to 6 days... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 19 | passed
[5 empty] | Executed | 57 | passed
Systems | passed | UTICA-172 - The hearing aid shall provide a battery life of greater than or equal to 4 days... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 20 | passed
[5 empty] | Executed | 58 | passed
Systems | passed | MASTER-17475 - Physical actuation of the user control to the point where tactile feedback is fe... | MASTER-47042 - AV005 - Automated Pushbutton Performance | CIC DVT 1 Run 1 - CIC SyRS Requirements Run 1 | Executed | 32 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 36 | passed
|  |  | MASTER-34397 - SV004 - Manual User Control Performance | CIC DVT 1 Run 1 - CIC SyRS Requirements Run 1 | Executed | 14 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-17577 - The hearing aid shall be capable of receiving audio from an installed loop syste... | MASTER-26390 - SV075 - Loop System | PPBTE DVT 1 Run 1 - PPBTE SyRS Requirements Run 1 | Executed | 11 | passed
[5 empty] | Executed | 13 | passed
Systems | passed | MASTER-19838 - The tinnitus masking device shall be capable of generating a broadband noise wit... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 41 | passed
Systems | passed | MASTER-19839 - The frequency response of the tinnitus masking device noise shall be adjustable... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19840 - The level of the tinnitus masking device noise shall be manually adjustable thro... | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 130 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 106 | passed
[5 empty] | Executed | 106 | passed
|  |  | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 137 | passed
[5 empty] | Executed | 247 | passed
Systems | passed | MASTER-19841 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19842 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 35 | passed
Systems | passed | MASTER-19866 - The hearing aid shall have a group delay of no longer than 5 ms. Note: A group d... | MASTER-28147 - SV080 - Group Delay | CIC DVT 1 Run 1 - CIC SyRS Requirements Run 1 | Executed | 10 | passed
Systems | passed | MASTER-48451 - The hearing aid shall produce no audible artifacts that are deemed objectionable... | MASTER-16449 - SV001 - General Audio Artifact | PPBTE IP5 - PPBTE SyRS Requirements Run 2 | Executed | 55 | passed
Systems | passed | UTICA-202 - If the requirement for no audible artifacts can't be achieved, the artifact(s) g... | MASTER-16449 - SV001 - General Audio Artifact | PPBTE IP5 - PPBTE SyRS Requirements Run 2 | Executed | 55 | passed
Systems | passed | MASTER-19678 - When Best Fit for a Flat 60 hearing loss, the hearing aid shall meet the perform... | MASTER-34339 - AV085 - Output Noise Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 41 | passed
[5 empty] | Executed | 73 | passed
|  |  | MASTER-34340 - AV084 - Output Noise Measurement | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
[5 empty] | Executed | 84 | passed
Systems | passed | UTICA-25729 - The hearing aid output noise shall not exceed the output noise limits defined in... | MASTER-34340 - AV084 - Output Noise Measurement | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
[5 empty] | Executed | 84 | passed
Systems | passed | UTICA-25733 - The hearing aid output noise shall not exceed the output noise limits defined in... | MASTER-34339 - AV085 - Output Noise Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 41 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-20000 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-20001 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
Systems | passed | MASTER-48455 - The hearing aid shall be capable of communicating with compatible Starkey access... | MASTER-49135 - SV234 - MFA Use Cases | Model S IP 6-5 | Executed | 163 | passed
[5 empty] | Executed | 166 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 172 | passed
|  |  | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 36 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-20002 - With no line-of-sight from the BLE Central to the user's head due to body-blocki... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-20011 - The hearing aid shall receive commands from the Remote Control at a distance of... | MASTER-34405 - SV117 - Remote Control Performance | CIC DVT 1 Run 1 - CIC SyRS Requirements Run 1 | Executed | 37 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 94 | passed
Systems | passed | MASTER-20064 - The hearing aid shall be programmable using the Starkey 2.4 GHz Programmer. | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 8 | passed
Systems | passed | MASTER-20065 - The hearing aid shall be programmable using the NOAHLink Wireless Programmer. No... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 8 | passed
ElectroAcoustics | passed | MASTER-17131 - The hearing aid shall use the NP5 platform. | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-42251 - The hearing aid shall include a Bluetooth radio. | UTICA-2760 - Verify by Inspection | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-42254 - The hearing aid shall include an inertial measurement unit. | UTICA-2760 - Verify by Inspection | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
|  |  | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 5 | passed
ElectroAcoustics | passed | MASTER-17164 - Noise resulting from hair rubbing against the hearing aid case and microphone co... | UTICA-2761 - Sound Quality- General / Audio Artifacts | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-17195 - When the earhook is installed in the fully seated position, the connection betwe... | UTICA-2759 - Sound Quality- Stability | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-17308 - The hearing aid shall be compatible with zinc air batteries. | UTICA-2757 - Zinc-air battery | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-17316 - The hearing aid shall be capable of electrical operation using a battery from al... | UTICA-3249 - Zinc-air battery | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-2757 - Zinc-air battery | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-17383 - The hearing aid shall operate at all battery voltages from 1.1 to 1.5 VDC, with... | UTICA-3249 - Zinc-air battery | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
|  |  | UTICA-2757 - Zinc-air battery | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-17498 - The hearing aid shall have one or more microphones, allowing the hearing aid to... | UTICA-2760 - Verify by Inspection | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-17503 - Hearing aid Directivity Index performance shall degrade no greater than 1 dB fro... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-17537 - The microphone cover, in conjunction with the overall hearing aid design, shall... | UTICA-2761 - Sound Quality- General / Audio Artifacts | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-59956 - The hearing aid shall detect the magnetic field from a telephone when held to th... | UTICA-3247 - Magnetic Switch | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-2762 - Magnetic Switch | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-17573 - The hearing aid shall be capable of receiving audio from a telephone handset via... | UTICA-2758 - Telecoil | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-17575 - The hearing aid shall be capable of receiving audio from a personal loop system... | UTICA-2758 - Telecoil | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-17577 - The hearing aid shall be capable of receiving audio from an installed loop syste... | UTICA-2758 - Telecoil | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | UTICA-2495 - OSPL90: Each hearing aid listed in TABLE1 - PEAK OSPL90 AND PEAK GAIN PER MATRIX... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | UTICA-2496 - ACOUSTIC GAIN: Each hearing aid listed in TABLE1 - PEAK OSPL90 AND PEAK GAIN PER... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-19808 - The hearing aid in-situ, occluded, frequency response shall be as flat as possib... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
|  |  | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-19790 - Design Guideline: The following illustration shows the desired frequency respons... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-59993 - The hearing aid shall comply with the Nordic homologation requirements defined... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 5 | passed
[5 empty] | Executed | 6 | passed
[5 empty] | Executed | 7 | passed
[5 empty] | Executed | 8 | passed
[5 empty] | Executed | 9 | passed
|  |  | UTICA-2758 - Telecoil | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-59991 - The hearing aid shall comply with the Australian homologation requirements defi... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 10 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 12 | passed
[5 empty] | Executed | 13 | passed
|  |  | UTICA-2758 - Telecoil | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 10 | passed
ElectroAcoustics | passed | MASTER-59992 - The hearing aid shall comply with the German homologation requirements defined i... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 14 | passed
ElectroAcoustics | passed | MASTER-19844 - POLARITY: The hearing aid shall preserve the polarity of the audio signal betwee... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4133 - Verify by Analysis | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19848 - When the Directionality Mode is Static Directional, the hearing aid shall have a... | MASTER-45252 - Directional- 3D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19849 - When in the HF Directional state and null steering is disabled, the hearing aid... | MASTER-45252 - Directional- 3D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-19853 - When the Directionality Mode is Static Directional, the hearing aid frequency re... | MASTER-45251 - Directional- DGC | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19854 - From 2000 Hz to 4000 Hz, the hearing aid frequency response in a HF Directional... | MASTER-45251 - Directional- DGC | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-19859 - When the full directional state and omnidirectional state are enabled, the follo... | MASTER-45249 - Directional-2D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19861 - When the full directional state and omnidirectional state are enabled, the follo... | MASTER-45249 - Directional-2D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-19862 - When the Directionality Mode is Adaptive Directionality, the following audio sti... | MASTER-45249 - Directional-2D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-19863 - When the Directionality Mode is Adaptive Directionality, the following audio sti... | MASTER-45249 - Directional-2D | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-48451 - The hearing aid shall produce no audible artifacts that are deemed objectionable... | UTICA-4122 - Sound Quality- General / Audio Artifacts | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-2761 - Sound Quality- General / Audio Artifacts | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | UTICA-202 - If the requirement for no audible artifacts can't be achieved, the artifact(s) g... |  |  | Executed | 2 | passed
ElectroAcoustics | passed | UTICA-156 - The hearing aid shall have a minimum lifespan of 3 years under normal use. | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | UTICA-157 - The hearing aid shall have a minimum lifespan of 5 years under normal use. | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 5 | passed
ElectroAcoustics | passed | MASTER-20084 - The materials used in the hearing aid shall not exceed 0.1 percent by weight in... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 6 | passed
ElectroAcoustics | passed | MASTER-20085 - The materials used in the hearing aid shall not exceed 0.01 percent by weight in... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 7 | passed
|  |  | UTICA-4133 - Verify by Analysis | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-60001 - The hearing aid shall comply with IEC 60601-1-2 for the following: * Radiated em... | UTICA-2764 - EMC- ESD | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-2765 - EMC- External Testing | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-60000 - The hearing aid shall comply with ETSI EN 301 489-17 for the following: * Radiat... | UTICA-2764 - EMC- ESD | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-2765 - EMC- External Testing | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-20102 - The hearing aid shall meet or exceed Category M2 and T2 (when configured with a... | MASTER-45261 - EMC- RF Immunity | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-4128 - EMC- RF Immunity | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-20104 - The hearing aid shall meet or exceed the following: * User Compatibility for Omn... | MASTER-45261 - EMC- RF Immunity | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4128 - EMC- RF Immunity | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
Mechanical | passed | MASTER-17102 - The hearing aid shall use Starkey's Sapphire design language. | UTICA-18158 - Industrial Design: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17115 - The hearing aid shall be available in the following standard colors: * Champagne... | UTICA-18158 - Industrial Design: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17107 - The hearing aid shall implement the Oslo 900 MHz Sapphire Power Plus BTE 13 hear... | UTICA-18158 - Industrial Design: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-60009 - The same earhook shall be used on hearing aids with the same industrial design.... | UTICA-18159 - Earhook: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17194 - The hearing aid shall allow the patient to rotate the earhook a minimum of +/- 1... | UTICA-18159 - Earhook: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17196 - The connection between the hearing aid and the earhook shall maintain an atmosph... | UTICA-19074 - Earhook: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17197 - The earhook shall maintain its position when placed in any position within its r... | UTICA-18159 - Earhook: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17199 - The connection between the hearing aid and the earhook shall withstand 10 cycles... | UTICA-19074 - Earhook: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17200 - The earhook shall retain a standard earmold tube with a removal force of 1.35 +/... | UTICA-19074 - Earhook: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17201 - The earhook shall be made of a clear finish material. | UTICA-19456 - Earhook: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17203 - The earhook shall not discolor for a period of at least one year. | UTICA-19456 - Earhook: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17204 - The earhook shall exhibit no more than minimal discoloration for a period of thr... | UTICA-19456 - Earhook: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-60013 - The same thin tube shall be used on hearing aids with the same industrial design... | UTICA-19079 - Thin Tube: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17222 - The connection between the hearing aid and the thin tube shall withstand 10 cycl... | UTICA-19080 - Thin Tube: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17223 - The connection between the hearing aid and the thin tube shall maintain an atmos... | UTICA-19080 - Thin Tube: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17224 - The thin tube header shall retain the tube with a removal force of greater than... | UTICA-19080 - Thin Tube: Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17225 - The thin tube shall be made of a clear material. | UTICA-19458 - Thin Tube: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17226 - The thin tube shall not discolor for a period of at least one year. | UTICA-19458 - Thin Tube: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-60016 - The thin tube shall exhibit no more than minimal discoloration for a period of t... | UTICA-19458 - Thin Tube: Material Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17268 - The hearing aid shall be compatible with the earbuds defined in the table below:... | UTICA-19083 - Earbud: Compatibility Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-59973 - The hearing aid design shall allow the induction coil to function with the telep... | UTICA-19087 - Right/Left Indicator: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17281 - The hearing aid design shall preclude the installation of the Right/Left indicat... | UTICA-19087 - Right/Left Indicator: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17282 - The hearing aid shall prohibit moisture ingress into the case at the mounting si... | UTICA-19088 - Right/Left Indicator: Assembly Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17283 - The Right/Left Indicator shall withstand all of the hearing aid's environmental... | UTICA-19088 - Right/Left Indicator: Assembly Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17285 - The Right/Left Indicator shall be no smaller than 1mm in diameter if round or 1m... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-19087 - Right/Left Indicator: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17289 - The hearing aid shall meet the other requirements specified herein without havin... | UTICA-19088 - Right/Left Indicator: Assembly Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17298 - The hearing aid shall use a size 13 battery. | UTICA-19461 - Battery: Size & Compatibility Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17313 - The hearing aid shall be capable of mechanically accommodating a battery from al... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-19461 - Battery: Size & Compatibility Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17386 - The battery door shall be a "C" style with no less than 80 degrees of the circle... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17387 - The battery door shall provide full access to the edge of the battery in the ope... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17389 - The battery door shall be labeled with a "+" sign (no quotes) that is no smaller... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 5 | passed
|  |  | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17390 - The battery door shall orient the battery such that the positive terminal is on... |  |  | Executed | 4 | passed
Mechanical | passed | MASTER-17391 - The battery door shall open in the direction of the hearing aid inner curve towa... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 5 | passed
Mechanical | passed | MASTER-17398 - The hearing aid shall have a detent position for the battery door which is no gr... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 6 | passed
Mechanical | passed | MASTER-17399 - The battery door shall open with a force in the range of 0.3 to 0.6 lbf. Note: T... | UTICA-19467 - Battery Door: Force Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17400 - When opening the battery door, the force to overcome the on/off detent shall be... | UTICA-19467 - Battery Door: Force Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17401 - When the battery door is between the closed and on/off detents, the force to mov... | UTICA-19467 - Battery Door: Force Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17403 - For all battery door positions from the on/off detent position and greater, the... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 7 | passed
Mechanical | passed | MASTER-17405 - For all battery door positions less than the on/off detent position, the hearing... | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 8 | passed
Mechanical | passed | MASTER-17410 - With the battery door in any position including the fully open position and the... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 9 | passed
Mechanical | passed | MASTER-17412 - When the battery door is fully open, it shall position the battery entirely outs... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 4 | passed
|  |  | UTICA-19856 - Battery Door: Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17419 - When the battery is installed with the incorrect orientation, the battery door s... |  |  | Executed | 2 | passed
Mechanical | passed | MASTER-17420 - Attempting to close the battery door with the battery in the wrong orientation s... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 3 | passed
|  |  | UTICA-19856 - Battery Door: Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17423 - The hearing aid shall limit moisture ingress into the case through the battery c... |  |  | Executed | 4 | passed
Mechanical | passed | MASTER-17425 - With a battery installed, the battery door and its mechanism of attachment to th... | UTICA-16706 - Battery Door Cycling | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-22627 - Battery Door: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17428 - The hearing aid shall power on and off as specified herein when subjected to the... | UTICA-16716 - Durability Testing | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-22627 - Battery Door: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17452 - The battery contacts shall be located such that they do not cover more than 40%... | UTICA-22630 - Battery Contacts: CAD Analysis/Physical Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17453 - With a battery installed, the battery contacts shall withstand exposure to, and... | UTICA-22633 - Battery Contacts: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17518 - The hearing aid design shall include protection to minimize foreign material ing... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 7 | passed
|  |  | UTICA-23260 - Microphone Protection: CAD Analysis/Physical Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | UTICA-364 - The hearing aid shall implement the Oslo 900 MHz Sapphire Power Plus BTE 13 hear... |  |  | Executed | 2 | passed
Mechanical | passed | MASTER-17525 - If a microphone cover is used, 15 cycles of removal and reinstallation of the co... | UTICA-23262 - Mic Cover: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19880 - The hearing aid shall be fully operable after exposure to storage temperatures r... | UTICA-16712 - Storage | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23257 - Hi/Lo Storage Temp Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19889 - With the battery door closed, the hearing aid shall provide ingress protection i... | UTICA-23241 - Dust Test - IP6x | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
|  |  | UTICA-23242 - Immersion Test - IPx8 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | UTICA-18505 - The hearing aid shall be fully operable after exposure to the drop shock environ... | UTICA-23239 - Drop Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19897 - The hearing aid exposure to the drop shock environment defined in Starkey specif... | UTICA-23239 - Drop Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-19906 - The hearing aid shall be fully operable during and after exposure to the light e... | UTICA-16711 - Light Sensitivity | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23237 - Light Sensitivity Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19907 - The hearing aid shall be fully operable after exposure to the heat and humidity... | UTICA-16710 - Damp Head Steady State | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23235 - DHSS/3yr & 5yr Lifespan Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19908 - The hearing aid shall be fully operable after exposure to the salt fog environme... | UTICA-23233 - Salt Fog Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19909 - The hearing aid exposure to the salt fog environment defined in Starkey specific... | UTICA-16709 - Salt Fog | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-23233 - Salt Fog Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-19912 - The hearing aid shall maintain its acoustic performance during and after exposur... | UTICA-23231 - Twist Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19916 - Plastic components shall meet the UV exposure test criteria defined in Starkey S... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23248 - UV Exposure Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19917 - Plastic components that are painted or have another decorative coating shall mee... | UTICA-23249 - Paint Adhesion Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19918 - Plastic components shall meet the alcohol soak test criteria defined in Starkey... | UTICA-23250 - Alcohol Soak Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19919 - Plastic components that are painted or have another decorative coating shall mee... | UTICA-23251 - Alcohol Rub Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19927 - The hearing aid parts shall meet the placement, alignment, color and quality cri... | UTICA-16708 - Internal Workmanship | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23224 - Final Visual Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19928 - The piece parts and assembly of the hearing aid shall meet the internal workmans... | UTICA-16708 - Internal Workmanship | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-23225 - Internal Workmanship | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | UTICA-156 - The hearing aid shall have a minimum lifespan of 3 years under normal use. | UTICA-23235 - DHSS/3yr & 5yr Lifespan Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | UTICA-157 - The hearing aid shall have a minimum lifespan of 5 years under normal use. | UTICA-23235 - DHSS/3yr & 5yr Lifespan Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-20084 - The materials used in the hearing aid shall not exceed 0.1 percent by weight in... | UTICA-23221 - European Directive 2011/65/EU (RoHS) - BTE13 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-20085 - The materials used in the hearing aid shall not exceed 0.01 percent by weight in... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23221 - European Directive 2011/65/EU (RoHS) - BTE13 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-20210 - All labeling on the hearing aid shall remain legible after exposure to the envir... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23219 - Label Permanence & Abrasion Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-46245 - The packaging shall provide protection such that exposure to the shipping distri... | UTICA-22444 - Packaging | 0000 - Utica CIC Packaging Testing | Executed | 1 | passed
|  |  | UTICA-23194 - Shipping & Packaging (BTE13): ASTM D4169 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
Quality | passed | MASTER-20078 - All hearing aid parts shall be latex-free. | MASTER-49164 - Latex / Phthalates / California Prop 65 | Utica Quality Test - Utica Pwr+ BTE | Executed | 1 | passed
Quality | passed | MASTER-20079 - The hearing aid shall be free of the following phthalates: * DEHP * DBP * BBP *... | MASTER-49164 - Latex / Phthalates / California Prop 65 | Utica Quality Test - Utica Pwr+ BTE | Executed | 2 | passed
Quality | passed | MASTER-20082 - Any hearing aid part that comes into contact with the patient's skin during norm... | MASTER-39834 - Plastic Parts Evaluation Procedure | Utica Quality Test - Utica Pwr+ BTE | Executed | 5 | passed
Product Management | passed | MASTER-60010 - The hearing aid shall be capable of being configured with adult and pediatric ea... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 1 | passed
Product Management | passed | MASTER-17193 - The hearing aid shall allow the professional to replace the earhook without usin... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 2 | passed
Product Management | passed | MASTER-17198 - Throughout the rotational range, the surface of the earhook shall align with the... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 3 | passed
Product Management | passed | MASTER-60015 - The hearing aid shall be capable of being configured with thin tubes. | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 4 | passed
Product Management | passed | MASTER-60014 - The thin tubes shall be available in left and right configurations with sizes co... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 5 | passed
Product Management | passed | MASTER-17220 - The hearing aid shall allow the professional to replace the thin tube without us... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 6 | passed
Product Management | passed | MASTER-17221 - The surface of the thin tube shall align with the surface of the case at the poi... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 7 | passed
Product Management | passed | MASTER-51866 - BTE earmolds shall be available in the style, tube type and material configurati... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 8 | passed
Product Management | passed | MASTER-51867 - BTE earmolds shall be available in the colors shown in the following table. Colo... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 9 | passed
Product Management | passed | MASTER-51870 - Earmolds shall have a Right/Left Indicator that is red to indicate right or blue... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 10 | passed
Product Management | passed | MASTER-17275 - The hearing aid shall have a Right/Left Indicator. The professional and the pati... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 11 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 1 | passed
Product Management | passed | MASTER-17280 - The Right/Left Indicator shall be field replaceable by the professional without... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 12 | passed
Product Management | passed | MASTER-17287 - The Right/Left Indicator shall be visible on the exterior of the hearing aid wit... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 13 | passed
Product Management | passed | MASTER-17288 - The Right/Left Indicator shall not be visible when the hearing aid is worn in th... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 14 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 3 | passed
Product Management | passed | MASTER-17388 - The battery door shall allow the patient to remove and replace the battery witho... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 15 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 5 | passed
Product Management | passed | MASTER-17393 - The battery door shall allow the patient to open and close the door without usin... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 16 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 6 | passed
Product Management | passed | MASTER-17394 - The battery door shall allow the patient to open and close the door without usin... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 17 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 7 | passed
Product Management | passed | MASTER-17397 - The hearing aid shall incorporate design cues to enable identification of the ac... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 18 | passed
Product Management | passed | MASTER-17411 - The battery door shall maintain its position, with and without a battery install... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 19 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 8 | passed
Product Management | passed | MASTER-17414 - The battery door or its attachment mechanism shall provide tactile feedback to i... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 20 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 9 | passed
Product Management | passed | MASTER-17415 - The battery door or its attachment mechanism shall provide a hard stop to indica... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 21 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 10 | passed
Product Management | passed | MASTER-59986 - The hearing aid shall have two user controls which are activated one at a time b... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 22 | passed
Product Management | passed | MASTER-17469 - The user control shall be higher than and have a different shape relative to the... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 23 | passed
Product Management | passed | MASTER-17471 - Physical actuation of the user control shall provide tactile feedback to the use... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 24 | passed
Product Management | passed | MASTER-17474 - Physical actuation of the user control shall not be audible to a normal hearing... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 25 | passed
Product Management | passed | MASTER-17476 - Physical shaking of the hearing aid shall not be audible to a normal hearing per... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 26 | passed
Product Management | passed | MASTER-17520 - If a microphone cover is used, it shall be field replaceable without the use of... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 27 | passed
Product Management | passed | MASTER-17524 - If a microphone cover is used, it shall provide tactile and audible feedback whe... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 28 | passed
Product Management | passed | MASTER-19926 - The hearing aid shall have no sharp edges or pointed protrusions that cut, poke,... | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 13 | passed
Regulatory | passed | MASTER-60001 - The hearing aid shall comply with IEC 60601-1-2 for the following: * Radiated em... | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 1 | passed
Regulatory | passed | MASTER-60000 - The hearing aid shall comply with ETSI EN 301 489-17 for the following: * Radiat... | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 2 | passed
Regulatory | passed | MASTER-20136 - The hearing aid shall comply with IEC 60601-2-66. | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 3 | passed
Regulatory | passed | MASTER-48458 - During normal operation in a no-fault condition, the temperature of the exterior... | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 4 | passed
Regulatory | passed | MASTER-41125 - During normal operation in a no-fault condition, the temperature of the exterior... | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 5 | passed
Regulatory | passed | MASTER-48459 - If the exterior surface of the hearing aid case can be greater than 41 degrees C... | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 6 | passed
Regulatory | passed | MASTER-41126 - When in a single fault condition, the temperature of the exterior surface of the... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 9 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 7 | passed
Regulatory | passed | MASTER-20138 - The hearing aid shall meet the specific absorption rate (SAR) requirements for r... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 10 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 8 | passed
Regulatory | passed | MASTER-20143 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 11 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 9 | passed
Regulatory | passed | MASTER-20144 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 12 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 10 | passed
Regulatory | passed | MASTER-20145 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 13 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 11 | passed
Regulatory | passed | MASTER-20147 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 14 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 12 | passed
Regulatory | passed | MASTER-20148 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 15 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 13 | passed
Regulatory | passed | MASTER-20152 - The hearing aid shall comply with ERC/REC 70-03, Annex 3, frequency band A. | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 16 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 14 | passed
Regulatory | passed | MASTER-20153 - The hearing aid shall comply with Commission Decision 2009/381/EC for operation... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 17 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 15 | passed
Regulatory | passed | MASTER-20154 - The hearing aid shall comply with ETSI EN 300 328 for operation within the band... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 18 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 16 | passed
Regulatory | passed | MASTER-20158 - The hearing aid shall comply with the requirements of RSS-Gen that are applicabl... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 19 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 17 | passed
Regulatory | passed | MASTER-20159 - The hearing aid shall comply with RSS-247 for operation within the band of 2400-... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 20 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 18 | passed
Regulatory | passed | MASTER-20160 - The hearing aid shall comply with RSS-102 for operation within the band of 2400-... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 21 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 19 | passed
Regulatory | passed | MASTER-20161 - The hearing aid shall comply with RSP-100. | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 22 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 20 | passed
Regulatory | passed | MASTER-20164 - The hearing aid shall comply with the performance and methods of measurement of... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 23 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 21 | passed
Regulatory | passed | MASTER-20167 - The hearing aid shall comply with KN 301 489-1. Note: Testing is performed using... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 24 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 22 | passed
Regulatory | passed | MASTER-42289 - The hearing aid shall comply with KN 301 489-17. Note: Testing is performed usin... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 25 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 23 | passed
Regulatory | passed | MASTER-42290 - The hearing aid shall comply with Korean standard KN 300 328 for operation withi... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 26 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 24 | passed
Regulatory | passed | MASTER-20170 - The hearing aid shall comply with China Ministry of Industry and Information Tec... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 27 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 25 | passed
Regulatory | passed | MASTER-20172 - The hearing aid shall comply with Japan Association of Radio Industries and Busi... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 28 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 26 | passed
Regulatory | passed | MASTER-20174 - The hearing aid shall be qualified to the standards defined by the Bluetooth Spe... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 29 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 27 | passed
Regulatory | passed | MASTER-50185 - The hearing aid shall be labeled as specified in Starkey specification DOCS-19-1... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 30 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 28 | passed
Regulatory | passed | MASTER-20222 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 31 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 29 | passed
Regulatory | passed | MASTER-20223 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 32 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 30 | passed
Regulatory | passed | MASTER-20224 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 33 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 31 | passed
Regulatory | passed | MASTER-20225 - The Tinnitus Therapy feature labeling for the patient shall include appropriate... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 34 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 32 | passed
Regulatory | passed | MASTER-20226 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 35 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 33 | passed
Regulatory | passed | MASTER-20227 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 36 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 34 | passed
Regulatory | passed | MASTER-20228 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 37 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 35 | passed
Regulatory | passed | MASTER-20229 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 38 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 36 | passed
Regulatory | passed | MASTER-20230 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 39 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 37 | passed
Regulatory | passed | MASTER-20231 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 40 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 38 | passed
Regulatory | passed | MASTER-20232 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 41 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 39 | passed
Regulatory | passed | MASTER-20233 - The Tinnitus Therapy feature labeling shall include information on possible misu... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 42 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 40 | passed
Regulatory | passed | MASTER-20234 - The Tinnitus Therapy feature labeling shall include information on basic operati... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 43 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 41 | passed
Component Engineering | passed | MASTER-41127 - When in a single fault condition, the temperature, in degrees C, of components w... | UTICA-6881 - Component Engineer Allocations | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
Design For Manufacturing | passed | MASTER-17654 - The hearing aid shall incorporate hydrophobic advanced coating (HAC) to mitigate... | UTICA-24618 - HAC process verification | Utica Power Plus BTE - DFM verification | Executed | N/A | passed
Design For Manufacturing | passed | MASTER-48453 - The hearing aid circuit shall provide programming pads that are accessible prior... | UTICA-24619 - Programming Pad Inspection | Utica CIC test case - Programming Pad Inspection | Executed | N/A | passed