## Sheet1
Utica Features on Paris HA Feature Metrics Report Grand Avenue Document ID: GQ-RPT-01039 Revision: 2.0
Description of data in this document: Function: This column designates which functional team is responsible for verifying the corresponding Systems requirement. Requirement Result: This column indicates the cumulative (across final test runs) verification result for the corresponding requirement. This is the cumulative result of the corresponding Test Resullt (Column H) data. Requirement: This column lists the requirement title as shown in Polarion and provides a hyperlink to the requirement in Polarion. Test Case: This column indicates which test case(s) contain steps that verify the requirement. Please note that Polarion includes all executed tests that reference the relevant requirement. Please also note that Polarion only includes verification results in this report from the most recently executed Test Run(s). For these reasons, some results included in this report may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Run: This column lists the test run(s) that contained the test case(s) (Column D) that verify the requirement. Please note that Polarion only includes verification results in this report from the most recently executed Test Run(s), which may or may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Result Status: This column indicates the execution status for each instance of the requirement in all executed test cases. Test Step: This column indicates the location of every instance of the requirement in the corresponding Test Case (Column D). Zero-step test cases (test cases that directly pass/fail an individual requirement) will indicate N/A in this column. Test Result: This column indicates the test result for every instance of the requirement in the corresponding Test Case (Column D). This column is the input for the Requirement Result (Column B) column.
Function | Requirement Result | Requirement | Test Case | Test Run | Test Result Status | Test Step | Test Result
Wireless | passed | MASTER-64684 - The hearing aid shall comply with MFi Accessory Interface Specification R33 Hear... | MASTER-64680 - MFi Certification | U31 - MFi Self Cert LEA1 | Executed | 2 | passed
Firmware | passed | MASTER-17676 - The hearing aid shall have 16 WOLA channels, allocated as follows: WOLA Channel... | MASTER-54271 - MECO (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-17679 - The hearing aid shall have 24 MECO channels, allocated as follows: MECO Channel... | MASTER-54271 - MECO (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-49262 - When the Double Tap sensitivity is set to Nominal, the hearing aid shall recogni... | MASTER-54286 - Basic Double Tap Functionality Test Procedure | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 11 | passed
Firmware | passed | MASTER-49260 - When the Double Tap sensitivity is set to Low, the hearing aid shall recognize a... | MASTER-54286 - Basic Double Tap Functionality Test Procedure | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 19 | passed
Firmware | passed | MASTER-50794 - When Double Tap is enabled, the hearing aid shall have the ability to detect a D... | MASTER-54286 - Basic Double Tap Functionality Test Procedure | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 24 | passed
Firmware | passed | MASTER-60726 - When the Trigger user control function is used to control the Personal Voice Ass... | MASTER-60850 - Basic Functionality Test with Android, ASHA Supported Phone | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 8 | passed
|  |  | MASTER-59028 - Basic Functionality Test with iPhone | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 9 | passed
Firmware | passed | MASTER-18070 - Upon receipt of a 2.4 GHz Remote Control command for any of the following items:... | MASTER-54255 - Automated Remote Control Verification (Nightly) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51043 - Upon receipt of a 2.4 GHz Remote Control command for any of the following items:... | MASTER-59813 - Two-way LEA2 streaming Takes Precedence over WiCROS Streaming with AutoStreaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 11 | passed
|  |  | MASTER-59529 - ASHA Streaming takes precedence over WiCROS with Auto Streaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 7 | passed
[5 empty] | Executed | 19 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 36 | passed
Firmware | passed | MASTER-18076 - Design Guideline: When an audio stream and microphone signal are being presented... | MASTER-54730 - Mic Volume Control from iPhone while in accessory streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 10 | passed
Firmware | passed | MASTER-18111 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-54481 - 2.4 GHz Remote Control - Special Features Button - TruLink | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-54255 - Automated Remote Control Verification (Nightly) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-18112 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 9 | passed
Firmware | passed | MASTER-18118 - The hearing aid shall allow Volume Reset on Memory Change to be enabled and disa... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 35 | passed
Firmware | passed | MASTER-18119 - For each channel within each memory location, the hearing aid shall allow a Maxi... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 34 | passed
Firmware | passed | MASTER-18167 - When a fitting session is active, upon a memory change, the hearing aid shall se... | MASTER-54291 - Speech MECO NP5 | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50018 - For each channel within each memory location, the hearing aid shall allow a Maxi... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 38 | passed
Firmware | passed | MASTER-50017 - The hearing aid shall allow a Volume Step Size of one of the following to be ass... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 39 | passed
Firmware | passed | MASTER-50022 - For each accessory stream input type, the hearing aid shall allow a default Acce... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 40 | passed
Firmware | passed | MASTER-18200 - The hearing aid shall allow a paired BLE Central to assign one Memory Environmen... | MASTER-54251 - Automated BLE Radio Verification (Nightly) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-18227 - The Tinnitus Therapy feature shall be designed with a common interface for all c... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 33 | passed
Firmware | passed | MASTER-18235 - The hearing aid shall be capable of generating and presenting a Tinnitus Therapy... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 15 | passed
Firmware | passed | MASTER-18259 - Upon receipt of a request to increase the Stimulus Level when the Stimulus is no... | MASTER-54813 - Disable Tinnitus Stimulus Fluctuation timeout | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18267 - Upon receipt of a request to decrease the Stimulus Level when the Stimulus is no... | MASTER-54789 - Tinnitus Interaction with Mute, UI, and E2E | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18271 - Upon expiration of the Disable Stimulus Fluctuation timeout, the hearing aid sha... | MASTER-54813 - Disable Tinnitus Stimulus Fluctuation timeout | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18272 - If an additional request to adjust the Stimulus Level is received prior to the D... | MASTER-54813 - Disable Tinnitus Stimulus Fluctuation timeout | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18295 - For each memory location, the hearing aid shall utilize the Fluctuation Magnitud... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 30 | passed
Firmware | passed | MASTER-18296 - For each memory location, the hearing aid shall utilize the Fluctuation Minimum... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 29 | passed
Firmware | passed | MASTER-18299 - When Tinnitus Therapy Stimulus fluctuation is enabled, the hearing aid shall cha... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18300 - When Tinnitus Therapy Stimulus fluctuation is enabled, the hearing aid shall cha... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-18301 - When Tinnitus Therapy Stimulus fluctuation is enabled, the hearing aid shall ens... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-18303 - When Tinnitus Therapy Stimulus fluctuation is enabled, the hearing aid shall shi... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
Firmware | passed | MASTER-18304 - When Tinnitus Therapy Stimulus fluctuation is enabled and the left edge of the p... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-18305 - When Tinnitus Therapy Stimulus fluctuation is enabled and the left edge of the p... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 6 | passed
Firmware | passed | MASTER-18306 - When Tinnitus Therapy Stimulus fluctuation is enabled, and the left edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 7 | passed
Firmware | passed | MASTER-18307 - When Tinnitus Therapy Stimulus fluctuation is enabled, and the left edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 8 | passed
Firmware | passed | MASTER-18308 - When Tinnitus Therapy Stimulus fluctuation is enabled, and the left edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 9 | passed
Firmware | passed | MASTER-18309 - When Tinnitus Therapy Stimulus fluctuation is enabled and the right edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 10 | passed
Firmware | passed | MASTER-18310 - When Tinnitus Therapy Stimulus fluctuation is enabled and the right edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 11 | passed
Firmware | passed | MASTER-18311 - When Tinnitus Therapy Stimulus fluctuation is enabled, and the right edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 12 | passed
Firmware | passed | MASTER-18312 - When Tinnitus Therapy Stimulus fluctuation is enabled and the right edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 13 | passed
Firmware | passed | MASTER-18313 - When Tinnitus Therapy Stimulus fluctuation is enabled and the right edge of the... | MASTER-54269 - Tinnitus (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 14 | passed
Firmware | passed | MASTER-18315 - When Tinnitus Therapy is enabled and an audio stream is being presented, the hea... | MASTER-54503 - Tinnitus Interaction with MRM Streaming | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 5 | passed
[5 empty] | Executed | 27 | passed
|  |  | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 9 | passed
[5 empty] | Executed | 10 | passed
|  |  | MASTER-58962 - OrCam Release Verification Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 192 | passed
Firmware | passed | MASTER-18316 - When the Tinnitus Therapy Stimulus has been stopped due to presenting an audio s... | MASTER-54503 - Tinnitus Interaction with MRM Streaming | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 11 | passed
[5 empty] | Executed | 33 | passed
|  |  | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 10 | passed
|  |  | MASTER-58962 - OrCam Release Verification Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 195 | passed
Firmware | passed | MASTER-18351 - For each memory location, the hearing aid shall allow Music Adaptation to be ena... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 22 | passed
Firmware | passed | MASTER-18352 - For each memory location, the hearing aid shall allow Music Adaptation E2E Coord... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 21 | passed
Firmware | passed | MASTER-18353 - For each memory location, the hearing aid shall allow the degree of Music adapta... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 20 | passed
Firmware | passed | MASTER-18354 - For each memory location, the hearing aid shall allow Machine Adaptation to be e... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 28 | passed
Firmware | passed | MASTER-18355 - For each memory location, the hearing aid shall allow Machine Adaptation E2E Coo... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 19 | passed
Firmware | passed | MASTER-18356 - For each memory location, the hearing aid shall allow the maximum gain reduction... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 18 | passed
Firmware | passed | MASTER-18357 - For each WOLA channel, the hearing aid shall allow the input level above which M... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 17 | passed
Firmware | passed | MASTER-18358 - The hearing aid shall allow the Machine adaptation time constants to be set as f... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 27 | passed
Firmware | passed | MASTER-18359 - For each memory location, the hearing aid shall allow Wind Adaptation to be enab... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 26 | passed
Firmware | passed | MASTER-18360 - For each memory location, the hearing aid shall allow Wind Adaptation E2E Coordi... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 25 | passed
Firmware | passed | MASTER-18361 - For each memory location, the hearing aid shall allow the maximum gain reduction... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 24 | passed
Firmware | passed | MASTER-18362 - For WOLA channels 1, 2 and 3, the hearing aid shall allow the input level above... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 16 | passed
Firmware | passed | MASTER-18363 - The hearing aid shall allow the Wind adaptation time constants to be set as foll... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 31 | passed
Firmware | passed | MASTER-18367 - The hearing aid shall classify the acoustic environment as one or two environmen... | MASTER-54945 - Advanced Environment Classification (AEC) Acoustic Test | HA-UTICA-04_27_2021 - Utica CIC Testing | Executed | 1 | passed
Firmware | passed | MASTER-18368 - When the broadband input sound level is less than 50 dB SPL, the hearing aid sha... | MASTER-54916 - AEC Quiet detection in Different Microphone Modes | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18369 - The hearing aid shall only allow the classification of the acoustic environment... | MASTER-54952 - AEC Wind Detection and Adaptation | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18370 - When the hearing aid does not classify the acoustic environment as Music, Machin... | MASTER-54941 - AEC Other Noise with and without Wind in Different Microphone Modes - High Frequency Directionality | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18373 - When presented with the Baseline Classifier Input, the hearing aid shall classif... | MASTER-54945 - Advanced Environment Classification (AEC) Acoustic Test | HA-UTICA-04_27_2021 - Utica CIC Testing | Executed | 1 | passed
Firmware | passed | MASTER-18379 - The hearing aid shall apply up to two environmental adaptations based on the cla... | MASTER-54945 - Advanced Environment Classification (AEC) Acoustic Test | HA-UTICA-04_27_2021 - Utica CIC Testing | Executed | 1 | passed
Firmware | passed | MASTER-18382 - The hearing aid shall request adaptation for a target environment only after pos... | MASTER-54273 - AEC (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18383 - The hearing aid shall request adaptation for a target environment only after pos... | MASTER-54945 - Advanced Environment Classification (AEC) Acoustic Test | HA-UTICA-04_27_2021 - Utica CIC Testing | Executed | 1 | passed
Firmware | passed | MASTER-18386 - When Music adaptation is enabled, the hearing aid shall apply Music adaptation i... | MASTER-54945 - Advanced Environment Classification (AEC) Acoustic Test | HA-UTICA-04_27_2021 - Utica CIC Testing | Executed | 1 | passed
Firmware | passed | MASTER-18388 - The hearing aid shall apply the following settings during Music adaptation: * Di... | MASTER-54984 - AEC Music Adaptation Verification WiCROS Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 3 | passed
[5 empty] | Executed | 7 | passed
|  |  | MASTER-54903 - AEC Music Adaptation Verification | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18389 - Upon exiting Music adaptation, the hearing aid shall apply Directionality State... | MASTER-54984 - AEC Music Adaptation Verification WiCROS Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 3 | passed
[5 empty] | Executed | 7 | passed
|  |  | MASTER-54903 - AEC Music Adaptation Verification | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18393 - When Machine adaptation is enabled, the hearing aid shall apply Machine adaptati... | MASTER-54930 - E2E Machine Noise Adaptation | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-55003 - E2E Machine Noise Adaptation WiCROS Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-18394 - When Machine adaptation is active, the hearing aid shall apply the Machine Adapt... | MASTER-54930 - E2E Machine Noise Adaptation | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-55003 - E2E Machine Noise Adaptation WiCROS Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-18397 - When Wind adaptation is enabled, the hearing aid shall apply Wind adaptation in... | MASTER-54952 - AEC Wind Detection and Adaptation | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18398 - When Wind Adaptation is active, the hearing aid shall apply the Wind Adaptation... | MASTER-54952 - AEC Wind Detection and Adaptation | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18399 - When the directionality state is Full Directionality or HF Directionality, wind... | MASTER-54916 - AEC Quiet detection in Different Microphone Modes | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18404 - When more than one gain reduction is simultaneously requested by Machine adaptat... | MASTER-54273 - AEC (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50224 - When a Binaural Noise Reduction stream is active, the hearing aid and the partne... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50225 - The hearing aid shall begin activation of the Binaural Noise Reduction stream wh... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50227 - Once the activation criteria have been met, the hearing aid shall complete Binau... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50229 - The hearing aid shall begin deactivating the Binaural Noise Reduction stream whe... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50228 - The hearing aid shall complete deactivating the Binaural Noise Reduction stream... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50230 - The hearing aid shall immediately cease the Binaural Noise Reduction stream when... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-50242 - When Monaural Noise Reduction is enabled and Binaural Noise Reduction is not act... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51891 - The hearing aid shall recognize a transient input signal when the amplitude of t... | MASTER-54270 - TNR (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51893 - When Transient Noise Reduction is enabled, no E2E streaming is active and a tran... | MASTER-54270 - TNR (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51894 - When Transient Noise Reduction is enabled, no E2E streaming is active and a tran... | MASTER-55030 - Dynamic BNR and TNR (Faraday) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51895 - Transient Noise Reduction shall attenuate the hearing aid output, applying the m... | MASTER-54270 - TNR (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51897 - The hearing aid shall apply Transient Noise Reduction gain reduction in addition... | MASTER-54270 - TNR (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-51898 - The hearing aid shall detect and attenuate transients for both microphone and te... | MASTER-54270 - TNR (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51899 - The hearing aid shall not process audible indicators through the Transient Noise... | MASTER-54263 - Voice Indicator DUT (Automated) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18557 - When Directionality Plus is enabled, the hearing aid shall apply MNR noise reduc... | MASTER-54921 - Directionality Plus Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18469 - For each MECO channel, the hearing aid shall allow the Directional Full-On Gain... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 10 | passed
Firmware | passed | MASTER-18474 - When in Full Directional State with Null Steering enabled, the hearing aid shall... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18477 - Upon memory change, the hearing aid shall switch to an initial directional state... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18478 - If a memory with null steering disabled is switched to a memory with null steeri... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18481 - When E2E Directionality is enabled, the hearing aid shall provide an indication... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
|  |  | MASTER-55356 - Interaction of WiCROS Streaming and Static Directionality | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-18482 - The hearing aid shall begin transition to Omnidirectional State within the Time... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18483 - The hearing aid shall begin transition to HF Directional state within the Time C... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18484 - The hearing aid shall begin transition to Full Directional state within the Time... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18486 - When a directional state transition is in progress, the hearing aid shall comple... | MASTER-54951 - Directionality Switching Logic Test (with Memory Transition, E2E / without Rear Microphone Input at all Noise Reduction Modes) | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18547 - When Null Steering is enabled, the hearing aid shall detect rear-hemisphere spee... | MASTER-54267 - Directionality (Automated Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18550 - For each WOLA channel other than the lowest channel, when Null Steering is enabl... | MASTER-54267 - Directionality (Automated Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-18554 - For each WOLA channel other than the lowest channel, when Null Steering is enabl... | MASTER-54267 - Directionality (Automated Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-18614 - When Null Steering is enabled, OPMAR is active, and Null Steering Away From Spee... | MASTER-54267 - Directionality (Automated Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
Firmware | passed | MASTER-18615 - When NULL Steering is enabled, the hearing aid shall determine the null steering... | MASTER-54267 - Directionality (Automated Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-18622 - When Noise Control is On, the hearing aid shall set its directionality state to... | MASTER-54321 - Noise Control - Force Mic Mode to Static Directional | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-18634 - For each memory location and MECO channel, the hearing aid shall allow the items... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 11 | passed
Firmware | passed | MASTER-53031 - When FBC is enabled, the hearing aid shall apply feedback cancellation to MECO c... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
[5 empty] | Executed | 5 | passed
Firmware | passed | MASTER-53034 - When FBC has been set to Adaptive - High Sensitivity, the hearing aid shall shif... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53035 - When FBC has been set to Adaptive - Low Sensitivity, the hearing aid shall shift... | MASTER-74721 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53036 - When FBC has been set to Static, the hearing aid shall not apply OPM to the hear... | MASTER-74721 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54306 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-53038 - The hearing aid shall apply OPM in MECO channels 5 through 24 when the energy in... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53039 - The hearing aid shall apply OPM in MECO channels 6 through 24 when either of the... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53040 - The hearing aid shall apply OPM in MECO channels 8 through 24 when the energy in... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-18700 - The hearing aid shall allow the Telephone Hold Time to be set in the range of at... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 23 | passed
Firmware | passed | MASTER-51332 - When the hearing aid receives a request to clear the fall event data, the hearin... | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55031 - Fall Alert Auto Detection and PDL | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 15 | passed
Firmware | passed | MASTER-51353 - When the hearing aid receives a request to clear the manual alert data, the hear... |  |  | Executed | 16 | passed
Firmware | passed | MASTER-60917 - Upon request, the hearing aid shall send logged On-Demand Adaptive Tuning featur... | MASTER-61018 - Adaptive Tuning Automatic Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
[5 empty] | Executed | 2 | passed
Firmware | passed | MASTER-60768 - Speech indicators shall be played through the Auto-Streaming memory MECO. | MASTER-54896 - Indicator Voice Level Dependency - OCL and fitting session test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19158 - The hearing aid shall complete initialization within 5 seconds of the applicatio... | MASTER-54827 - Power On Delay | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-54989 - HA initialization completed within 5 seconds of the application of power to the device | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-19610 - For each memory location, the hearing aid shall utilize the Splitting Frequency... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19611 - For each memory location, the hearing aid shall utilize the Knee Frequency that... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19612 - For each memory location, the hearing aid shall utilize the Warping Ratio that h... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19613 - For each memory location, the hearing aid shall utilize the Frequency Translatio... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19617 - When Frequency Translation is enabled, the hearing aid shall output no audio in... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19618 - When Frequency Translation is enabled, the hearing aid shall translate up to two... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19619 - When Frequency Translation is enabled, the hearing aid shall move a translated p... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19620 - When Frequency Translation is enabled, the hearing aid shall increase the gain o... | MASTER-54268 - Frequency Translation (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19621 - The hearing aid shall not process tone indicators through the Frequency Translat... | MASTER-54888 - Voice Indicator - FT | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-54263 - Voice Indicator DUT (Automated) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19639 - The hearing aid shall allow the Auto Acclimatization Time Per Step to be set wit... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-19644 - The hearing aid shall begin counting time spent on a subsequent step in the Auto... | MASTER-56894 - AAM Time Tracking with Data Log Automation Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19648 - The hearing aid shall use the same method of determination and tracking of time... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
|  |  | MASTER-56894 - AAM Time Tracking with Data Log Automation Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56285 - The hearing aid shall allow a value to be set indicating that the Receiver Cable... | MASTER-55028 - Manual Tests of Heart Rate Measurement GATT Attributes | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-19657 - The hearing aid shall be capable of uniquely identifying the following resistor... | MASTER-54882 - RID Manual Test When in Fitting | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-20052 - The hearing aid shall use G.722 decoding at 64 Kbps for the digital transmission... | MASTER-53261 - Verification Method of "Analysis" or "Inspection" | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | UTICA-5360 - The hearing aid shall use the "Faraday - 0x3A, Hopper - 0x82, Wired SSI Communic... | MASTER-64568 - Verify Correct Wireless and Wired SSI versions | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-50670 - Upon receipt of a pairing deletion command for one of the following accessory de... | MASTER-54828 - Delete accessory pairing information | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 17 | passed
|  |  | MASTER-55550 - MRM and HA Pairing Deletion | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
Firmware | passed | MASTER-59912 - When an audio stream is being received from an iOS Device, and the iOS Device re... | MASTER-58793 - LEA 2.0 Streaming Test(Calls) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 5 | passed
[5 empty] | Executed | 17 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 33 | passed
Firmware | passed | MASTER-59915 - The hearing aid shall stop transmitting the microphone stream upon request from... | MASTER-58793 - LEA 2.0 Streaming Test(Calls) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 9 | passed
[5 empty] | Executed | 11 | passed
Firmware | passed | MASTER-59918 - When transmitting a microphone stream to an iOS Device, and the corresponding in... | MASTER-58793 - LEA 2.0 Streaming Test(Calls) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 13 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 37 | passed
Firmware | passed | MASTER-19254 - When receiving and presenting an audio stream from an iOS Device, Android Device... | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
[5 empty] | Executed | 6 | passed
[5 empty] | Executed | 9 | passed
[5 empty] | Executed | 10 | passed
|  |  | MASTER-58964 - Tinnitus Interaction with LEA 2.0 Streaming with Auto Streaming on | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 7 | passed
[5 empty] | Executed | 13 | passed
[5 empty] | Executed | 21 | passed
|  |  | MASTER-59642 - FBC Mode During iPhone Streaming when MicAdjustment is 100 | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 6 | passed
[5 empty] | Executed | 9 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 20 | passed
|  |  | MASTER-59479 - FBC Mode in ASHA Streaming when AutoStreaming is On | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 6 | passed
[5 empty] | Executed | 8 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 20 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
|  |  | MASTER-58963 - Tinnitus Interaction with ASHA Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 7 | passed
[5 empty] | Executed | 16 | passed
|  |  | MASTER-59484 - FBC Mode During LEA2 Bidrectional Phone Call with AutoStreaming ON | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 6 | passed
[5 empty] | Executed | 8 | passed
[5 empty] | Executed | 10 | passed
[5 empty] | Executed | 17 | passed
[5 empty] | Executed | 19 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 33 | passed
Firmware | passed | MASTER-50080 - When receiving and presenting an audio stream from an iOS Device or an accessory... | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 21 | passed
|  |  | MASTER-56160 - Mic Attenuation During Streaming and Auto-Streaming | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-58955 - OrCam Interaction with Mic Volume | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 27 | passed
|  |  | MASTER-59744 - Mic Attenuation During and After ASHA Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
[5 empty] | Executed | 13 | passed
|  |  | MASTER-59631 - Mic Attenuation During and After iPhone Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 5 | passed
[5 empty] | Executed | 9 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 20 | passed
|  |  | MASTER-54506 - MRM Interaction with Mic Volume | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 19 | passed
Firmware | passed | MASTER-19255 - When an audio stream from an accessory device other than an OrCam Visual Assista... | MASTER-56160 - Mic Attenuation During Streaming and Auto-Streaming | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-58964 - Tinnitus Interaction with LEA 2.0 Streaming with Auto Streaming on | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 11 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 23 | passed
|  |  | MASTER-59930 - OrCam Stream Loss, Stream Ending, and Stream Stop Test | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 156 | passed
|  |  | MASTER-59742 - Mic Volume After iPhone Streaming with Volume Reset On and Off | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 6 | passed
[5 empty] | Executed | 12 | passed
|  |  | MASTER-59379 - Connection loss between a Hearing Aid and an Android Phone During ASHA Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 6 | passed
[5 empty] | Executed | 10 | passed
|  |  | MASTER-58963 - Tinnitus Interaction with ASHA Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 11 | passed
[5 empty] | Executed | 18 | passed
Firmware | passed | MASTER-63412 - When an audio stream from an iOS Device or Android Device ends, the hearing aid... | MASTER-64371 - Verify the use of StrmEndMicVolumeState for LEA and Android Streams | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
[5 empty] | Executed | 2 | passed
Firmware | passed | MASTER-42177 - Upon receipt of a Start Accessory Stream Request initiated via user control or 2... | MASTER-54495 - MFA Mic Interaction with MRM | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 26 | passed
|  |  | MASTER-58431 - Nuance and HAs Pair within 3 min | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 6 | passed
[5 empty] | Executed | 34 | passed
|  |  | MASTER-58792 - ASHA Multiple Connection and Streaming Priority Test with AutoStreaming On | HA-UticaCont_9_13_2021 - Utica Continuation Hopper RSL10 Adoption and LEA2 | Executed | 16 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 27 | passed
|  |  | MASTER-54563 - Verify Start/Stop Stream indicator from TV Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 4 | passed
|  |  | MASTER-54744 - MFA and HAs Pair within 3 min | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 10 | passed
|  |  | MASTER-54746 - Multiple Connection Interaction Test | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-42191 - Upon receipt of a request to terminate audio streaming from a specific paired ac... | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 15 | passed
|  |  | MASTER-54754 - Directed Scan and Stop Stream via GATT with Multiple Devices Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 23 | passed
Firmware | passed | MASTER-19317 - When Auto-Streaming is enabled, an audio stream is received, and the hearing aid... | MASTER-54807 - Stream Boost Interaction Between E2E Phone Streaming and Accessory Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 4 | passed
[5 empty] | Executed | 8 | passed
[5 empty] | Executed | 23 | passed
|  |  | MASTER-54501 - PDL and MRM Streaming with Auto Streaming ON | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 5 | passed
|  |  | MASTER-54508 - iPhone Streaming Interaction with MRM Streaming | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
[5 empty] | Executed | 25 | passed
|  |  | MASTER-54818 - Interaction Between iPhone Streaming and Accessory Streaming | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 4 | passed
[5 empty] | Executed | 21 | passed
|  |  | MASTER-54716 - Push Button UI and Remote Control During TV streaming with Streaming Boost ON. | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-54491 - E2E Phone Streaming and MRM Interaction (Auto Streaming ON) | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 22 | passed
|  |  | MASTER-54492 - Push Button E2E UI and Remote Control During MRM Streaming with Auto Streaming ON | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 27 | passed
|  |  | MASTER-54264 - Monoaural Automated TV Streamer & Hearing Aid Interaction Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
|  |  | MASTER-55792 - Interaction Between iPhone Streaming and MFA Phone Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 4 | passed
[5 empty] | Executed | 18 | passed
|  |  | MASTER-58962 - OrCam Release Verification Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 78 | passed
Firmware | passed | UTICA-6875 - When Auto-Streaming is enabled and Auto-Streaming settings are received from the... | MASTER-54508 - iPhone Streaming Interaction with MRM Streaming | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 25 | passed
Firmware | passed | MASTER-19355 - The hearing aid shall include its current memory location in the message initiat... | MASTER-55888 - E2E Phone Streaming With ATR and E2E UI | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-19358 - For each memory location, the hearing aid shall have an E2E Phone Streaming Micr... | MASTER-54752 - Microphone Attenuation During E2E Phone Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-19370 - When E2E Phone Streaming is in progress, the presentation of the audio signal th... | MASTER-54750 - E2E Phone Streaming Delay Verification | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-19394 - When the current memory location is configured as CROS, the WiCROS Partner heari... | MASTER-54986 - TV Streaming Takes Precedence over WiCROS | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 10 | passed
|  |  | MASTER-55350 - Interaction of ATR and WiCROS Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 3 | passed
|  |  | MASTER-55354 - Receiving a Ring Event while WiCROS Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 31 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 43 | passed
|  |  | MASTER-55356 - Interaction of WiCROS Streaming and Static Directionality | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
|  |  | MASTER-59813 - Two-way LEA2 streaming Takes Precedence over WiCROS Streaming with AutoStreaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 18 | passed
[5 empty] | Executed | 24 | passed
|  |  | MASTER-59529 - ASHA Streaming takes precedence over WiCROS with Auto Streaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 13 | passed
[5 empty] | Executed | 22 | passed
Firmware | passed | MASTER-19395 - When the current memory location is configured as BiCROS, the WiCROS Partner hea... | MASTER-54986 - TV Streaming Takes Precedence over WiCROS | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 33 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 76 | passed
|  |  | MASTER-55350 - Interaction of ATR and WiCROS Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
|  |  | MASTER-55354 - Receiving a Ring Event while WiCROS Streaming | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 44 | passed
[5 empty] | Executed | 54 | passed
|  |  | MASTER-55356 - Interaction of WiCROS Streaming and Static Directionality | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
|  |  | MASTER-59813 - Two-way LEA2 streaming Takes Precedence over WiCROS Streaming with AutoStreaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 24 | passed
|  |  | MASTER-59529 - ASHA Streaming takes precedence over WiCROS with Auto Streaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 22 | passed
Firmware | passed | MASTER-51115 - When the current memory location is configured as Hearing Aid Only, the WiCROS P... | MASTER-55350 - Interaction of ATR and WiCROS Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 12 | passed
[5 empty] | Executed | 27 | passed
|  |  | MASTER-59813 - Two-way LEA2 streaming Takes Precedence over WiCROS Streaming with AutoStreaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 3 | passed
|  |  | MASTER-59529 - ASHA Streaming takes precedence over WiCROS with Auto Streaming On | HA-UticaCont_8_30_2021 - Utica Continuation Faraday RSL10 Adoption and LEA2 | Executed | 2 | passed
Firmware | passed | MASTER-19403 - The WiCROS balance control shall adjust the WiCROS Partner hearing aid microphon... | MASTER-55010 - Balance Control Sprinkler - WiCros | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-19404 - Each adjustment of the WiCROS balance control shall result in equal but opposite... | MASTER-55010 - Balance Control Sprinkler - WiCros | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 8 | passed
[5 empty] | Executed | 23 | passed
Firmware | passed | MASTER-19406 - The WiCROS system shall only apply the WiCROS balance control offset to the micr... | MASTER-55350 - Interaction of ATR and WiCROS Streaming | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 35 | passed
[5 empty] | Executed | 42 | passed
Firmware | passed | MASTER-50193 - When the BLE Central requests Brain Score data, the hearing aid shall transmit t... | MASTER-54251 - Automated BLE Radio Verification (Nightly) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-49276 - The hearing aid shall have a method to enable and disable activity tracking. | MASTER-54253 - Activity Tracker Test Cases | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 8 | passed
Firmware | passed | MASTER-47585 - The hearing aid shall lose no more than 15 minutes worth of activity data per ea... | MASTER-54253 - Activity Tracker Test Cases | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
Firmware | passed | MASTER-47583 - The hearing aid shall log the most recent 8 hours' worth of activity data. | MASTER-54253 - Activity Tracker Test Cases | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-47590 - During a fitting session initiated by mobile application or fitting software, th... | MASTER-54253 - Activity Tracker Test Cases | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
[5 empty] | Executed | 6 | passed
[5 empty] | Executed | 7 | passed
Firmware | passed | MASTER-47589 - When the mobile application requests activity data, the hearing aid shall transm... | MASTER-54253 - Activity Tracker Test Cases | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51330 - When the hearing aid battery voltage is above the low voltage level, the hearing... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56914 - When the hearing aid battery voltage is above the low voltage level, the WiCROS... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51809 - The hearing aid shall have a method to enable or disable Manual Alert. | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56916 - The WiCROS Partner hearing aid shall have a method to enable or disable Manual A... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51808 - When Fall Detection is enabled, the hearing aid shall have a method to set the F... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51811 - The hearing aid shall detect a fall under the following sensitivity setting, fal... | MASTER-56166 - Fall Detection performance and sensitivity Verification | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51338 - The hearing aid shall detect at least 80% of fall events when using the lab fall... | MASTER-56166 - Fall Detection performance and sensitivity Verification | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51341 - The hearing aid shall detect a non-fall event (Activity of Daily Living) with at... | MASTER-56166 - Fall Detection performance and sensitivity Verification | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51815 - The hearing aid shall have the ability to generate alerts 10 seconds after: * th... | MASTER-55015 - Auto Fall Detection power on Behavior | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 8 | passed
[5 empty] | Executed | 13 | passed
Firmware | passed | MASTER-51817 - When the hearing aid detects a fall event or a Manual Alert is triggered, the he... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 6 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 9 | passed
[5 empty] | Executed | 17 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51377 - Upon receiving an alert sent receipt from BLE Central within 15 seconds of sendi... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 16 | passed
[5 empty] | Executed | 29 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 18 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-55032 - ASB Streaming will be stopped while Auto Fall or Manual Alert Impact occurred | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 7 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
[5 empty] | Executed | 4 | passed
[5 empty] | Executed | 8 | passed
Firmware | passed | MASTER-51378 - When an E2E present Alert Sent indicator command is received within 15 seconds o... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51820 - When an E2E present Alert Sent indicator command is received, after a maximum of... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51380 - If the hearing aid does not receive an alert sent receipt from BLE Central withi... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 8 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 12 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 35 | passed
Firmware | passed | MASTER-51822 - When an E2E restore the hearing aid user control to programmed settings and pres... | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 10 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 15 | passed
|  |  | MASTER-55032 - ASB Streaming will be stopped while Auto Fall or Manual Alert Impact occurred | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-51821 - When an E2E restore the hearing aid user control to programmed settings and pres... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 8 | passed
Firmware | passed | MASTER-51387 - Upon receiving an alert canceled request within 90 seconds of receiving the aler... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 32 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 21 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 30 | passed
|  |  | MASTER-55032 - ASB Streaming will be stopped while Auto Fall or Manual Alert Impact occurred | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 7 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-51388 - When an E2E alert canceled request command is received, the hearing aid shall tr... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 22 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
[5 empty] | Executed | 4 | passed
Firmware | passed | MASTER-51400 - Upon receiving an alert canceled receipt from BLE Central within 15 seconds of s... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55024 - verify Auto Fall Event Detection | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 30 | passed
Firmware | passed | MASTER-51401 - When an E2E alert is canceled and present Alert Canceled indicator command is re... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 32 | passed
|  |  | MASTER-55032 - ASB Streaming will be stopped while Auto Fall or Manual Alert Impact occurred | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 8 | passed
Firmware | passed | MASTER-51823 - When an E2E alert is canceled and present Alert Canceled indicator command is re... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51828 - If the hearing aid does not receive an alert canceled receipt from BLE Central w... | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51829 - When an E2E present the Alert Failure indicator command is received within 15 se... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51827 - When an E2E present the Alert Failure indicator command is received after a maxi... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51411 - If the hearing aid receives an alert delivered receipt from BLE Central, or afte... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55014 - Auto Fall detection interacts with NFMI (I2C Bus sharing) | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 21 | passed
[5 empty] | Executed | 29 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 10 | passed
Firmware | passed | MASTER-51409 - When an E2E alert delivered command is received, the hearing aid shall restore t... | MASTER-55016 - Binaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-51825 - If the hearing aid receives an alert received receipt from BLE Central, the hear... | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 10 | passed
Firmware | passed | MASTER-51426 - If the hearing aid is presenting an accessory audio stream when sending an alert... | MASTER-55032 - ASB Streaming will be stopped while Auto Fall or Manual Alert Impact occurred | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 2 | passed
[5 empty] | Executed | 7 | passed
Firmware | passed | MASTER-51423 - If the hearing aid is presenting an audio stream from BLE Central and when the h... | MASTER-55029 - Manual Fall Emergency Alert while E2E phone streaming is active or inactive | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 16 | passed
Firmware | passed | MASTER-51460 - Upon detecting a fall event, the hearing aid shall store the 3-axis acceleromete... | MASTER-55023 - Verify RAW IMU data after Auto Fall Event injection | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 8 | passed
Firmware | passed | MASTER-51468 - Upon receiving a request for the Fall Event Accelerometer Data, the hearing aid... | MASTER-55013 - Monaural Automated Fall Management Test | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-55019 - Verify fall events and timing sequence using light blue mobile application | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 6 | passed
Firmware | passed | MASTER-56552 - The hearing aid shall present the Self Check Test Signal as a sum of 5 sinusoida... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56550 - The hearing aid shall present the Self Check Test Signal for no greater than 2 s... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56553 - For each microphone input, the hearing aid shall determine the energy in WOLA ch... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56554 - When Self Check Baseline measurements exist and a subsequent Self Check Baseline... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56555 - Upon completion of Self Check Baseline, and the microphone input was greater tha... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56556 - Upon completion of Self Check Baseline, and the microphone input was less than -... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56559 - Upon receipt of a command to perform Initiated Self Check, the hearing aid shall... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56561 - For each microphone input, the hearing aid shall determine the energy in the WOL... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56562 - When at least one microphone has a grade of Pass or Questionable from the Initia... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56564 - Each hour, the hearing aid shall sample the input from the active microphone(s)... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 6 | passed
Firmware | passed | MASTER-56567 - When both microphones have a grade of Bad from the Initiated Self Check and Time... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 6 | passed
Firmware | passed | MASTER-56568 - When both microphones have a grade of Bad from the Initiated Self Check and Time... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 6 | passed
Firmware | passed | MASTER-56569 - Upon receipt of a command to perform Initiated Self Check, the hearing aid shall... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56571 - Upon completion of Initiated Self Check, the hearing aid shall return the result... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-56573 - Upon completion of Initiated Self Check, the hearing aid shall store the results... | MASTER-56052 - Self Check Baseline and Diagnostic Automated Test (SSI) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
|  |  | MASTER-56054 - Self Check Automated Test (Interface Check and In Fitting Session, low battery mode) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-60855 - Upon the receipt of a Start Stream command from the BLE Central, the hearing aid... | MASTER-60850 - Basic Functionality Test with Android, ASHA Supported Phone | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 9 | passed
|  |  | MASTER-59028 - Basic Functionality Test with iPhone | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 10 | passed
Firmware | passed | MASTER-56889 - When a subsequent request to initiate the Personal Voice Assistant feature is re... | MASTER-60844 - PVA Automated Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-60857 - When the audio stream from the hearing aid is configured as the input source, an... | MASTER-60844 - PVA Automated Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-60856 - When the audio stream from the hearing aid is configured as the input source, an... | MASTER-60844 - PVA Automated Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 4 | passed
Firmware | passed | MASTER-61844 - Upon the receipt of a Personal Voice Assistant activated response from the BLE C... | MASTER-60844 - PVA Automated Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-61843 - When the hearing aids discontinue Personal Voice Assistant mic upstream, if Adap... | MASTER-60844 - PVA Automated Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-57698 - When the hearing aid is in a Low Power State and is picked up and handled, the h... | MASTER-57110 - Auto Sleep/Wake Functionality (while Not in Fitting Session) Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 41 | passed
Firmware | passed | MASTER-60912 - Upon activation of On-Demand Adaptive Tuning, the hearing aid shall sample and c... | MASTER-61611 - ODAT - Wind Categorization | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
|  |  | MASTER-61016 - ODAT Acoustic Environment Categorization | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-60910 - When On-Demand Adaptive Tuning has been deactivated due to initiation of an iOS,... | MASTER-61219 - ODAT, iPhone Streaming and Telecoil Memory, E2E Phone Streaming Interaction with AutoStreaming On Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
[5 empty] | Executed | 14 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 56 | passed
[5 empty] | Executed | 63 | passed
|  |  | MASTER-60658 - ODAT - LEA - Auto Streaming On and Off | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 7 | passed
Firmware | passed | MASTER-60913 - Upon receipt of a request to deactivate On-Demand Adaptive Tuning, the hearing a... |  |  | Executed | 12 | passed
Firmware | passed | MASTER-60915 - When On-Demand Adaptive Tuning is active, the hearing aid shall deactivate On-De... | MASTER-60658 - ODAT - LEA - Auto Streaming On and Off | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 3 | passed
Firmware | passed | MASTER-64596 - Upon activation of On-Demand Adaptive Tuning, the hearing aid shall sample and c... | MASTER-61016 - ODAT Acoustic Environment Categorization | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-64595 - On request from the mobile application, the hearing aid shall return information... | MASTER-61018 - Adaptive Tuning Automatic Tests | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 3 | passed
Systems | passed | MASTER-60681 - When configured with a "rocker switch" user control, the hearing aid shall suppo... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 51 | passed
Systems | passed | MASTER-60682 - When configured with a "rocker switch" user control and a hearing aid feature is... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 29 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 5 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 51 | passed
Systems | passed | MASTER-60679 - When configured with an IMU, the hearing aid shall support assignment of the fol... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 14 | passed
[5 empty] | Executed | 14 | passed
|  |  | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 39 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-60680 - When configured with an IMU, the WiCROS Partner hearing aid shall support assign... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 40 | passed
|  |  | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 64 | passed
[5 empty] | Executed | 103 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-60683 - When a binaural fitting session is initiated and the Push and Hold - Top Power O... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 44 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 6 | passed
Systems | passed | MASTER-60692 - The hearing aid shall recognize the following action as a Short Press/Release: A... |  |  | Executed | 11 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 41 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 50 | passed
[5 empty] | Executed | 54 | passed
Systems | passed | MASTER-60693 - The hearing aid shall recognize the following action as an Extended Press: A pre... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 12 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 47 | passed
[5 empty] | Executed | 49 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 126 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 139 | passed
Systems | passed | MASTER-60690 - The hearing aid shall recognize the following actions as an Extended Press/Hold:... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 9 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 51 | passed
[5 empty] | Executed | 55 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 164 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 177 | passed
Systems | passed | MASTER-60691 - The hearing aid shall recognize the following action as an Extended Press/Releas... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 10 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 47 | passed
[5 empty] | Executed | 49 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 126 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 139 | passed
Systems | passed | MASTER-49262 - When the Double Tap sensitivity is set to Nominal, the hearing aid shall recogni... | MASTER-48456 - SV232 - Double Tap Detection Performance Test | IP7 - General IP7 Testing | Executed | 45 | passed
Systems | passed | MASTER-49260 - When the Double Tap sensitivity is set to Low, the hearing aid shall recognize a... | MASTER-48456 - SV232 - Double Tap Detection Performance Test | IP7 - General IP7 Testing | Executed | 45 | passed
Systems | passed | MASTER-49266 - The hearing aid shall detect a Double Tap with at least an 80% success rate. | MASTER-48456 - SV232 - Double Tap Detection Performance Test | IP7 - General IP7 Testing | Executed | 30 | passed
[5 empty] | Executed | 34 | passed
Systems | passed | MASTER-60704 - When the Increment user control function is used to control a hearing aid featur... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 14 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 36 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-60706 - When the Decrement user control function is used to control a hearing aid featur... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 15 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 39 | passed
[5 empty] | Executed | 49 | passed
Systems | passed | MASTER-60721 - When the Trigger user control function is used to control the Mute/Unmute featur... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 75 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 22 | passed
Systems | passed | MASTER-60722 - When the Trigger user control function is used to control the Mute/Unmute featur... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 77 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 23 | passed
Systems | passed | MASTER-60727 - When the Trigger user control function is used to control the Start/Stop Accesso... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 26 | passed
Systems | passed | MASTER-60728 - When the Trigger user control function is used to control the Start/Stop Accesso... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 31 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 27 | passed
Systems | passed | MASTER-60725 - When the Trigger user control function is used to control the Manual Alert featu... | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 19 | passed
[5 empty] | Executed | 62 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-60904 - When the Trigger user control function is used to control the On-Demand Adaptive... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 23 | passed
[5 empty] | Executed | 23 | passed
Systems | passed | MASTER-60903 - When the Trigger user control function is used to control the On-Demand Adaptive... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 24 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-60735 - When a hearing aid feature is assigned to a Short Press gesture-function, E2E is... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 30 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 135 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 141 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 182 | passed
[5 empty] | Executed | 187 | passed
[5 empty] | Executed | 205 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 148 | passed
[5 empty] | Executed | 152 | passed
[5 empty] | Executed | 210 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 218 | passed
[5 empty] | Executed | 223 | passed
[5 empty] | Executed | 241 | passed
Systems | passed | MASTER-60732 - When a hearing aid feature is assigned to a Short Press gesture-function, no feature is assigned to a Press and Hold gesture-function, E2E is enabled, and an Extended Press/Release user control action is recognized, the hearing aid shall synchronize the assigned feature with its E2E Partner. Note: When only a Short Press gesture-function is assigned, the Extended Press/Release user control action is also utilized to control the feature. The result is that whenever the user control is released the user control function will be activated. This is done as a convenience to the user. | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 28 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 92 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 207 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 243 | passed
Systems | passed | MASTER-60733 - When a hearing aid feature is assigned to a Push and Hold gesture-function and E... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 17 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 29 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 137 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 182 | passed
[5 empty] | Executed | 187 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 213 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 150 | passed
[5 empty] | Executed | 210 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 218 | passed
[5 empty] | Executed | 223 | passed
[5 empty] | Executed | 248 | passed
[5 empty] | Executed | 249 | passed
Systems | passed | MASTER-60737 - When a hearing aid feature is assigned to a Push and Hold gesture-function for t... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 31 | passed
Systems | passed | MASTER-18038 - When a BLE Central Remote Control command is received to change the microphone v... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 123 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 132 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 123 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 132 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 123 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 132 | passed
|  |  | MASTER-42121 - SV170 - Mobile Application - Native Application and Controls | IP5 - General IP5 Testing | Executed | 43 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 49 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 64 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 71 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 80 | passed
Systems | passed | MASTER-18039 - When a BLE Central Remote Control command is received to change the microphone v... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 84 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 87 | passed
Systems | passed | MASTER-18040 - When a BLE Central Remote Control command is received to mute the microphone vol... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 66 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 84 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 91 | passed
[5 empty] | Executed | 92 | passed
Systems | passed | MASTER-18041 - When a BLE Central Remote Control command is received to unmute the microphone v... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 86 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 91 | passed
[5 empty] | Executed | 92 | passed
Systems | passed | MASTER-18042 - The hearing aid shall receive and execute BLE Central Remote Control commands fo... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 32 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-51028 - The WiCROS system shall apply BLE Central Remote Control microphone volume comma... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 88 | passed
Systems | passed | MASTER-18044 - When the Remote Control Volume Up command is received, the hearing aid shall rec... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 48 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 74 | passed
[5 empty] | Executed | 75 | passed
Systems | passed | MASTER-18045 - When the Remote Control Volume Down command is received, the hearing aid shall r... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 48 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 74 | passed
[5 empty] | Executed | 75 | passed
Systems | passed | MASTER-19431 - When the WiCROS system receives a Remote Control Volume Up command, regardless o... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 20 | passed
Systems | passed | MASTER-19432 - When the WiCROS system receives a Remote Control Volume Down command, regardless... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 20 | passed
Systems | passed | MASTER-18050 - When a BLE Central Remote Control command is received to change to a specified m... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 35 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 47 | passed
|  |  | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 83 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 93 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 28 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18053 - When the hearing aid receives a 2.4 GHz Remote Control command to increment the... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 133 | passed
[5 empty] | Executed | 142 | passed
[5 empty] | Executed | 151 | passed
Systems | passed | MASTER-19429 - When the WiCROS system receives a Remote Control Multi-Memory Control command, r... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 17 | passed
Systems | passed | MASTER-18056 - When a BLE Central Remote Control command is received to change the tinnitus the... | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 124 | passed
[5 empty] | Executed | 124 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 132 | passed
[5 empty] | Executed | 134 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
[5 empty] | Executed | 140 | passed
[5 empty] | Executed | 142 | passed
[5 empty] | Executed | 147 | passed
[5 empty] | Executed | 151 | passed
[5 empty] | Executed | 155 | passed
Systems | passed | MASTER-18057 - When a BLE Central Remote Control command is received to change the tinnitus the... | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 127 | passed
[5 empty] | Executed | 127 | passed
Systems | passed | MASTER-18065 - Upon receipt of a BLE Central Remote Control Mute command and the hearing aid mi... | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 28 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 41 | passed
[5 empty] | Executed | 41 | passed
Systems | passed | MASTER-18066 - Upon receipt of a BLE Central Remote Control Mute command and the hearing aid mi... | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 52 | passed
[5 empty] | Executed | 70 | passed
[5 empty] | Executed | 52 | passed
[5 empty] | Executed | 70 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 59 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 59 | passed
[5 empty] | Executed | 62 | passed
Systems | passed | MASTER-51533 - Upon receipt of a BLE Central Remote Control command for any of the following it... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 26 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 29 | passed
Systems | passed | MASTER-51031 - Upon receipt of a BLE Central Remote Control Mute command and the WiCROS Partner... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 91 | passed
Systems | passed | MASTER-51032 - Upon receipt of a BLE Central Remote Control Mute command and the WiCROS Partner... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 93 | passed
[5 empty] | Executed | 99 | passed
Systems | passed | MASTER-51029 - Upon receipt of a BLE Central Remote Control WiCROS Mute command and the WiCROS... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 77 | passed
Systems | passed | MASTER-51030 - Upon receipt of a BLE Central Remote Control WiCROS Mute command, the WiCROS aud... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 96 | passed
Systems | passed | MASTER-51534 - Upon receipt of a BLE Central Remote Control command for any of the following it... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 101 | passed
Systems | passed | MASTER-18068 - The hearing aid shall recognize a received 2.4 GHz Remote Control Mute Toggle co... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 203 | passed
[5 empty] | Executed | 205 | passed
Systems | passed | MASTER-18069 - The hearing aid shall recognize a received 2.4 GHz Remote Control Mute Toggle co... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 206 | passed
[5 empty] | Executed | 208 | passed
Systems | passed | MASTER-18070 - Upon receipt of a 2.4 GHz Remote Control command for any of the following items:... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 163 | passed
[5 empty] | Executed | 166 | passed
Systems | passed | MASTER-19449 - The WiCROS system shall recognize a received 2.4 GHz Remote Control Mute Toggle... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 28 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
Systems | passed | MASTER-19450 - The WiCROS system shall recognize a received 2.4 GHz Remote Control Mute Toggle... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 30 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 38 | passed
Systems | passed | MASTER-51043 - Upon receipt of a 2.4 GHz Remote Control command for any of the following items:... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 44 | passed
Systems | passed | MASTER-18073 - When a BLE Central Remote Control command is received to change the BLE Central... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 26 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 26 | passed
Systems | passed | MASTER-18074 - When a BLE Central Remote Control command is received to change the BLE Central... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 28 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-50001 - When a BLE Central Remote Control command is received to change the BLE Central... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 38 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 67 | passed
Systems | passed | MASTER-50004 - When the Remote Control Volume Up command is received, the hearing aid shall rec... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 179 | passed
[5 empty] | Executed | 181 | passed
Systems | passed | MASTER-50005 - When the Remote Control Volume Down command is received, the hearing aid shall r... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 180 | passed
[5 empty] | Executed | 182 | passed
Systems | passed | MASTER-51662 - Upon receipt of an audio stream volume increase request from a 2.4 GHz streaming... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 65 | passed
Systems | passed | MASTER-51663 - Upon receipt of an audio stream volume decrease request from a 2.4 GHz streaming... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 71 | passed
Systems | passed | MASTER-18081 - When the 2.4 GHz Remote Control Home command is received, the hearing aid shall... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 208 | passed
Systems | passed | MASTER-19435 - When the WiCROS system receives a Home command from a remote control device, reg... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 23 | passed
Systems | passed | MASTER-42106 - The hearing aid shall allow one Special Feature from the following list to be as... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 26 | passed
Systems | passed | MASTER-18086 - When the hearing aid's Special Feature is set to None, it shall ignore any Remot... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 245 | passed
Systems | passed | MASTER-42110 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 262 | passed
[5 empty] | Executed | 274 | passed
[5 empty] | Executed | 278 | passed
Systems | passed | MASTER-42111 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 255 | passed
[5 empty] | Executed | 267 | passed
Systems | passed | MASTER-42112 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 257 | passed
[5 empty] | Executed | 269 | passed
Systems | passed | MASTER-51056 - The WiCROS system shall respond to a Special Feature command when the Special Fe... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 61 | passed
Systems | passed | MASTER-18101 - The hearing aid shall recognize a received BLE Central command as a Noise Contro... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 30 | passed
Systems | passed | MASTER-18102 - The hearing aid shall recognize a received BLE Central command as a Noise Contro... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
Systems | passed | MASTER-18104 - The hearing aid shall recognize a received 2.4 GHz Remote Control command as a N... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 236 | passed
Systems | passed | MASTER-18105 - The hearing aid shall recognize a received 2.4 GHz Remote Control command as a N... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 238 | passed
Systems | passed | MASTER-19535 - The WiCROS system shall respond to a Special Feature command when the Special Fe... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 54 | passed
Systems | passed | MASTER-18107 - The hearing aid shall activate Memory 4 and sound the appropriate indicator (Tel... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 216 | passed
Systems | passed | MASTER-18108 - The hearing aid shall activate the memory location that was active prior to Memo... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 218 | passed
Systems | passed | MASTER-18109 - The hearing aid shall remain in Memory 4 and sound the appropriate indicator (Te... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 220 | passed
Systems | passed | MASTER-19444 - The WiCROS system shall respond to a Special Feature command when the Special Fe... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 25 | passed
Systems | passed | MASTER-18111 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 225 | passed
Systems | passed | MASTER-18112 - When the following conditions are true: * The hearing aid receives a Remote Cont... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 227 | passed
Systems | passed | MASTER-18114 - When a memory change occurs, the hearing aid shall present the Tinnitus Therapy... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 234 | passed
[5 empty] | Executed | 236 | passed
[5 empty] | Executed | 244 | passed
Systems | passed | MASTER-19447 - The WiCROS system shall respond to a Special Feature command when the Special Fe... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 49 | passed
Systems | passed | MASTER-51076 - When a BLE Central Remote Control command is received to change the WiCROS Balan... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 86 | passed
Systems | passed | MASTER-18116 - When in a fitting session, the hearing aid shall implement changes to volume set... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 85 | passed
[5 empty] | Executed | 86 | passed
Systems | passed | MASTER-24072 - The hearing aid shall allow a volume range of 20 dB to be assigned out of a sele... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 27 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 22 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 23 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 36 | passed
[5 empty] | Executed | 185 | passed
Systems | passed | MASTER-64681 - The hearing aid shall have a method for assigning following Volume Step Size in... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 3 | passed
Systems | passed | MASTER-18144 - For each channel, upon receipt of a request via user control to increase its vol... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 46 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 41 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 50 | passed
[5 empty] | Executed | 54 | passed
Systems | passed | MASTER-42113 - When the volume level is "in-between" user control step levels after a BLE Centr... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 86 | passed
Systems | passed | MASTER-19459 - Upon receipt of a request to increase the volume level and the volume setting is... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 109 | passed
Systems | passed | MASTER-19460 - Upon receipt of a request to increase the volume level and the volume setting is... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 111 | passed
Systems | passed | MASTER-18152 - For each channel, upon receipt of a request via user control to decrease its vol... | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 39 | passed
Systems | passed | MASTER-42114 - When the volume level is "in-between" user control step levels after a BLE Centr... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 85 | passed
Systems | passed | MASTER-19462 - Upon receipt of a request to decrease the volume level and the volume setting is... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 112 | passed
Systems | passed | MASTER-19463 - Upon receipt of a request to decrease the volume level and the volume setting is... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 114 | passed
Systems | passed | MASTER-18160 - The hearing aid shall allow a level within the volume control range to be assign... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 34 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 29 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 30 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 39 | passed
Systems | passed | MASTER-18162 - Upon power-up, the hearing aid shall set its volume level to the Use Gain settin... | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 37 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 46 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 119 | passed
Systems | passed | MASTER-18163 - Upon a change in memory location, when Volume Reset on Memory Change is enabled,... |  |  | Executed | 17 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 41 | passed
[5 empty] | Executed | 67 | passed
[5 empty] | Executed | 69 | passed
Systems | passed | MASTER-18164 - Upon a change in memory location, when Volume Reset on Memory Change is disabled... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 28 | passed
[5 empty] | Executed | 52 | passed
[5 empty] | Executed | 78 | passed
Systems | passed | MASTER-18171 - When the hearing aid receives a request to change its volume level via user cont... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 283 | passed
[5 empty] | Executed | 285 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 24 | passed
[5 empty] | Executed | 29 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 63 | passed
[5 empty] | Executed | 65 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 79 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 176 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 92 | passed
[5 empty] | Executed | 210 | passed
[5 empty] | Executed | 212 | passed
Systems | passed | MASTER-50015 - When in a fitting session, the hearing aid shall implement changes to volume set... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 41 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-50020 - For each channel, upon receipt of a request via user control to increase its vol... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 30 | passed
[5 empty] | Executed | 55 | passed
|  |  | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 65 | passed
[5 empty] | Executed | 69 | passed
Systems | passed | MASTER-50019 - For each channel, upon receipt of a request via user control to decrease its vol... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 35 | passed
[5 empty] | Executed | 52 | passed
|  |  | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 71 | passed
Systems | passed | MASTER-50021 - When an accessory device audio stream starts or switches to another input type,... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 30 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 122 | passed
[5 empty] | Executed | 127 | passed
[5 empty] | Executed | 132 | passed
[5 empty] | Executed | 137 | passed
[5 empty] | Executed | 142 | passed
[5 empty] | Executed | 147 | passed
Systems | passed | MASTER-50013 - When an accessory device audio stream stops or the audio stream switches to anot... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 148 | passed
[5 empty] | Executed | 149 | passed
Systems | passed | MASTER-50012 - When the hearing aid receives a request to change its volume level via user cont... | MASTER-49114 - SV002G - User Control - Accessory Volume Control | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 28 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 52 | passed
[5 empty] | Executed | 55 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 172 | passed
Systems | passed | MASTER-18182 - The hearing aid shall have 6 memory locations: Memory 1, Memory 2, Memory 3, Mem... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 31 | passed
Systems | passed | MASTER-18189 - The hearing aid shall allow Memory 2, Memory 3 and Memory 4 to be individually e... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 3 | passed
Systems | passed | MASTER-18190 - The hearing aid shall allow Memory 5 and Memory 6 to be enabled and disabled. No... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 32 | passed
Systems | passed | MASTER-18200 - The hearing aid shall allow a paired BLE Central to assign one Memory Environmen... | MASTER-42122 - SV169 - Mobile Application - Custom Memories | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 28 | passed
Systems | passed | MASTER-18201 - The hearing aid shall allow one memory environment to be assigned to each memory... | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-18209 - When the hearing aid receives a request to decrement its memory location, it sha... | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 53 | passed
[5 empty] | Executed | 64 | passed
|  |  | MASTER-75313 - SV097 - DTMF | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 221 | passed
Systems | passed | MASTER-18203 - Upon power-up, the hearing aid shall activate Memory 1. | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
[5 empty] | Executed | 50 | passed
Systems | passed | MASTER-50183 - The hearing aid shall allow Memory 4 and Memory 5 to independently be included i... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 7 | passed
[5 empty] | Executed | 34 | passed
Systems | passed | MASTER-18208 - When the hearing aid receives a request to activate its next memory location, it... | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 51 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-18210 - While transitioning from one memory location to another, the hearing aid shall m... | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 85 | passed
Systems | passed | MASTER-18211 - The hearing aid shall activate one memory location at a time. | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 83 | passed
[5 empty] | Executed | 94 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 94 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 93 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-18212 - Upon activating a memory location, the hearing aid shall configure itself as pre... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 5 | passed
Systems | passed | UTICA-19406 - When the hearing aid transitions to Memory 1, 2, 3 or 4, and the following condi... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 108 | passed
Systems | passed | MASTER-18222 - When the hearing aid transitions to Memory 5 and memory indicators are enabled,... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 34 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 34 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 68 | passed
Systems | passed | MASTER-18223 - When the hearing aid transitions to Memory 6 via a user initiated memory change,... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 34 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 34 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
Systems | passed | MASTER-19454 - When the WiCROS system transitions to Memory 1, 2, 3 or 4, and Memory indicators... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 24 | passed
Systems | passed | MASTER-51078 - When the WiCROS system transitions to Memory 5 and memory indicators are enabled... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 104 | passed
Systems | passed | MASTER-18225 - When in a fitting session, the hearing aid shall implement changes to Tinnitus T... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 25 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 137 | passed
Systems | passed | MASTER-18229 - For each memory location, the hearing aid shall allow Tinnitus Therapy to be ena... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 10 | passed
Systems | passed | MASTER-18235 - The hearing aid shall be capable of generating and presenting a Tinnitus Therapy... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 69 | passed
[5 empty] | Executed | 69 | passed
Systems | passed | MASTER-18236 - The hearing aid shall generate and present the Tinnitus Therapy Stimulus indepen... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 382 | passed
Systems | passed | MASTER-18237 - The generation and presentation of the Tinnitus Therapy Stimulus shall have no e... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 383 | passed
Systems | passed | MASTER-18238 - The hearing aid function settings and processing shall have no effect on the gen... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 382 | passed
Systems | passed | MASTER-18239 - When any of the audible indicators is presented, the hearing aid shall only outp... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 324 | passed
Systems | passed | MASTER-18240 - When Tinnitus Therapy is enabled and fluctuation is disabled, the hearing aid sh... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-18242 - For each memory location, the hearing aid shall allow the Tinnitus Therapy Stimu... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-18244 - For each memory location, the hearing aid shall allow the Tinnitus Therapy Stimu... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-18248 - When Tinnitus Therapy Stimulus Level Control is enabled, the hearing aid shall p... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 135 | passed
[5 empty] | Executed | 168 | passed
[5 empty] | Executed | 135 | passed
[5 empty] | Executed | 168 | passed
Systems | passed | MASTER-18249 - The hearing aid shall utilize the Tinnitus Therapy Stimulus Level step size that... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 134 | passed
[5 empty] | Executed | 167 | passed
[5 empty] | Executed | 134 | passed
[5 empty] | Executed | 167 | passed
Systems | passed | MASTER-18254 - Upon receipt of a request via user control to increase the Stimulus Level when t... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 28 | passed
[5 empty] | Executed | 269 | passed
[5 empty] | Executed | 270 | passed
Systems | passed | MASTER-18255 - Upon receipt of a request to increase the Stimulus Level when the Stimulus is be... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 269 | passed
[5 empty] | Executed | 270 | passed
[5 empty] | Executed | 273 | passed
[5 empty] | Executed | 274 | passed
Systems | passed | MASTER-18262 - Upon receipt of a request via user control to decrease the Stimulus Level when t... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 276 | passed
[5 empty] | Executed | 277 | passed
Systems | passed | MASTER-18264 - Upon receipt of a request to decrease the Stimulus Level when the Stimulus is be... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 276 | passed
[5 empty] | Executed | 277 | passed
Systems | passed | MASTER-18275 - The hearing aid shall utilize the level within the adjustable range that has bee... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
|  |  | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 125 | passed
[5 empty] | Executed | 158 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 158 | passed
Systems | passed | MASTER-18276 - Upon power-up, the hearing aid shall set its Stimulus Level to the power-on Tinn... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 10 | passed
|  |  | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 125 | passed
[5 empty] | Executed | 158 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 158 | passed
Systems | passed | MASTER-18282 - Upon a change in memory location, when Volume Reset on Memory Change is enabled,... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 188 | passed
[5 empty] | Executed | 192 | passed
[5 empty] | Executed | 214 | passed
[5 empty] | Executed | 218 | passed
Systems | passed | MASTER-18283 - Upon a change in memory location, when Volume Reset on Memory Change is disabled... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 199 | passed
[5 empty] | Executed | 225 | passed
Systems | passed | MASTER-18284 - When a fitting session is active, upon a change in memory location, the hearing... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 180 | passed
[5 empty] | Executed | 207 | passed
Systems | passed | MASTER-18286 - When the hearing aid receives a request to change its Stimulus Level and the Tin... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 320 | passed
[5 empty] | Executed | 321 | passed
[5 empty] | Executed | 322 | passed
[5 empty] | Executed | 323 | passed
Systems | passed | MASTER-18287 - When the hearing aid receives a request via user control (including E2E commands... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
[5 empty] | Executed | 32 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 39 | passed
[5 empty] | Executed | 44 | passed
Systems | passed | MASTER-18288 - When the hearing aid receives a request to change its Stimulus Level and the Tin... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 137 | passed
[5 empty] | Executed | 139 | passed
Systems | passed | MASTER-18293 - For each memory location, the hearing aid shall allow Tinnitus Therapy Stimulus... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 251 | passed
Systems | passed | MASTER-18294 - For each memory location, the hearing aid shall utilize the Fluctuation Rate of... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 251 | passed
Systems | passed | MASTER-18297 - The hearing aid shall vary the level and frequency of the Tinnitus Therapy Stimu... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 266 | passed
Systems | passed | MASTER-18315 - When Tinnitus Therapy is enabled and an audio stream is being presented, the hea... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 242 | passed
Systems | passed | MASTER-18316 - When the Tinnitus Therapy Stimulus has been stopped due to presenting an audio s... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 77 | passed
Systems | passed | MASTER-19838 - The tinnitus masking device shall be capable of generating a broadband noise wit... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 41 | passed
Systems | passed | MASTER-19839 - The frequency response of the tinnitus masking device noise shall be adjustable... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19840 - The level of the tinnitus masking device noise shall be manually adjustable thro... | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 106 | passed
[5 empty] | Executed | 106 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 130 | passed
|  |  | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 137 | passed
[5 empty] | Executed | 247 | passed
Systems | passed | MASTER-19841 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19842 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 35 | passed
Systems | passed | MASTER-18321 - Upon receipt of a User Control Mute Request, the hearing aid shall sound the Mut... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 75 | passed
[5 empty] | Executed | 84 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 115 | passed
Systems | passed | MASTER-18322 - Upon receipt of a User Control Mute Request, when the Tinnitus Therapy Stimulus... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 90 | passed
Systems | passed | MASTER-42116 - Upon receipt of a User Control Mute Request, when the hearing aid microphone is... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 75 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-18324 - Upon receipt of a User Control Mute Request, the hearing aid shall mute the foll... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 75 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-19469 - Upon receipt of a WiCROS Transmitter User Control Mute Request, the WiCROS syste... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
Systems | passed | MASTER-51081 - Upon receipt of a WiCROS Transmitter User Control Mute Request, the WiCROS syste... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 46 | passed
Systems | passed | MASTER-51080 - Upon receipt of a WiCROS Partner hearing aid User Control Mute Request, the WiCR... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 49 | passed
[5 empty] | Executed | 50 | passed
Systems | passed | MASTER-18332 - Upon receipt of a User Control Unmute Request, the hearing aid shall sound the U... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 49 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 115 | passed
Systems | passed | MASTER-18333 - Upon receipt of a User Control Unmute Request, the hearing aid shall: * Set the... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 62 | passed
[5 empty] | Executed | 90 | passed
Systems | passed | MASTER-18334 - Upon receipt of a user control request (including E2E commands) for any of the f... | MASTER-26724 - SV002D - User Control - Mute | IP9 - General IP9 Testing & Regression | Executed | 77 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 107 | passed
|  |  | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 300 | passed
|  |  | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 178 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 80 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 191 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 227 | passed
Systems | passed | MASTER-19473 - Upon receipt of a user control request (including E2E commands) for any of the f... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 38 | passed
Systems | passed | MASTER-19474 - Upon receipt of a WiCROS Transmitter User Control Unmute Request, the WiCROS sys... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 47 | passed
Systems | passed | MASTER-51084 - Upon receipt of a WiCROS Transmitter User Control Unmute Request, the WiCROS sys... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 24 | passed
Systems | passed | MASTER-55395 - The WiCROS Transmitter shall use the following settings for Advanced Environment... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 147 | passed
Systems | passed | MASTER-18367 - The hearing aid shall classify the acoustic environment as one or two environmen... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 25 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-18368 - When the broadband input sound level is less than 50 dB SPL, the hearing aid sha... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
Systems | passed | MASTER-18369 - The hearing aid shall only allow the classification of the acoustic environment... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 19 | passed
[5 empty] | Executed | 19 | passed
Systems | passed | MASTER-18370 - When the hearing aid does not classify the acoustic environment as Music, Machin... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 72 | passed
Systems | passed | MASTER-18379 - The hearing aid shall apply up to two environmental adaptations based on the cla... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 25 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 25 | passed
Systems | passed | MASTER-18386 - When Music adaptation is enabled, the hearing aid shall apply Music adaptation i... | MASTER-34371 - SV142: Music AEC-Audio Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 18 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18388 - The hearing aid shall apply the following settings during Music adaptation: * Di... | MASTER-34371 - SV142: Music AEC-Audio Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 19 | passed
Systems | passed | MASTER-18389 - Upon exiting Music adaptation, the hearing aid shall apply Directionality State... | MASTER-34371 - SV142: Music AEC-Audio Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
Systems | passed | MASTER-18393 - When Machine adaptation is enabled, the hearing aid shall apply Machine adaptati... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 27 | passed
Systems | passed | MASTER-18394 - When Machine adaptation is active, the hearing aid shall apply the Machine Adapt... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 21 | passed
[5 empty] | Executed | 22 | passed
Systems | passed | MASTER-18397 - When Wind adaptation is enabled, the hearing aid shall apply Wind adaptation in... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 19 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 19 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-18398 - When Wind Adaptation is active, the hearing aid shall apply the Wind Adaptation... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 26 | passed
[5 empty] | Executed | 26 | passed
Systems | passed | MASTER-18399 - When the directionality state is Full Directionality or HF Directionality, wind... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-18404 - When more than one gain reduction is simultaneously requested by Machine adaptat... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-50225 - The hearing aid shall begin activation of the Binaural Noise Reduction stream wh... | MASTER-34643 - SV205 - BNR Switching and Feature Interaction Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 54 | passed
[5 empty] | Executed | 58 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 73 | passed
[5 empty] | Executed | 75 | passed
Systems | passed | MASTER-50227 - Once the activation criteria have been met, the hearing aid shall complete Binau... | MASTER-34643 - SV205 - BNR Switching and Feature Interaction Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
Systems | passed | MASTER-50229 - The hearing aid shall begin deactivating the Binaural Noise Reduction stream whe... | MASTER-34643 - SV205 - BNR Switching and Feature Interaction Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 25 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 56 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-50228 - The hearing aid shall complete deactivating the Binaural Noise Reduction stream... | MASTER-34643 - SV205 - BNR Switching and Feature Interaction Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 56 | passed
Systems | passed | MASTER-50230 - The hearing aid shall immediately cease the Binaural Noise Reduction stream when... | MASTER-34643 - SV205 - BNR Switching and Feature Interaction Test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 47 | passed
[5 empty] | Executed | 51 | passed
Systems | passed | UTICA-25181 - With BNR in the active state and a stationary white noise field is present and w... | MASTER-46617 - SV225 - BNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 51 | passed
Systems | passed | UTICA-25179 - With BNR in the active state and when an 8-talker babble input is present and wh... | MASTER-46617 - SV225 - BNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 70 | passed
Systems | passed | UTICA-25176 - With BNR in the active state and when presented with a 70 dB SPL mixture of MNR_... | MASTER-46617 - SV225 - BNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 83 | passed
Systems | passed | UTICA-25680 - With BNR in the active state and when the following audio stimuli are present, e... | MASTER-46617 - SV225 - BNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 92 | passed
Systems | passed | UTICA-25188 - When a stationary white noise field is present and when configured in the curren... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 29 | passed
Systems | passed | UTICA-25679 - When a restaurant noise input (see attached reference) is present and when confi... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 53 | passed
Systems | passed | UTICA-25190 - When presented with MNR_BNR_Speech_Input_1 and when configured in the current me... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 37 | passed
Systems | passed | UTICA-25191 - When presented with a 70 dB SPL mixture of MNR_BNR_Speech_Input_1 and 8-talker b... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 76 | passed
Systems | passed | MASTER-51773 - When MNR is enabled and the input level is less than 45 dB SPL, the hearing aid... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 120 | passed
Systems | passed | MASTER-51774 - When MNR is enabled and the input level is less than 45 dB SPL, the hearing aid... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 121 | passed
Systems | passed | MASTER-51777 - For the Autocoil, Telecoil and Loop memory environments, when the input level is... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 122 | passed
Systems | passed | MASTER-51778 - For the Music memory environment, when the input level is less than 45 dB SPL, t... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 123 | passed
Systems | passed | MASTER-51895 - Transient Noise Reduction shall attenuate the hearing aid output, applying the m... | MASTER-56423 - AV247 - Transient Noise Reduction (TNR) - Acoustic Input | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 36 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-51898 - The hearing aid shall detect and attenuate transients for both microphone and te... | MASTER-56423 - AV247 - Transient Noise Reduction (TNR) - Acoustic Input | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 36 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-18562 - For each memory location, the hearing aid shall allow Directionality Plus to be... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 135 | passed
Systems | passed | MASTER-18557 - When Directionality Plus is enabled, the hearing aid shall apply MNR noise reduc... | MASTER-62839 - SV099B - Speech in Noise - MNR (Paris and forward) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 36 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18463 - For each memory location, the hearing aid shall allow the fitting software to en... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 11 | passed
Systems | passed | MASTER-18464 - For each memory location, the hearing aid shall allow the fitting software to en... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 13 | passed
Systems | passed | MASTER-18465 - The hearing aid shall allow the fitting software to set the switching thresholds... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
Systems | passed | MASTER-18466 - The hearing aid shall allow the fitting software to set the directional switchin... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 127 | passed
[5 empty] | Executed | 148 | passed
[5 empty] | Executed | 150 | passed
[5 empty] | Executed | 160 | passed
[5 empty] | Executed | 162 | passed
[5 empty] | Executed | 170 | passed
[5 empty] | Executed | 172 | passed
[5 empty] | Executed | 182 | passed
[5 empty] | Executed | 184 | passed
[5 empty] | Executed | 192 | passed
[5 empty] | Executed | 194 | passed
[5 empty] | Executed | 204 | passed
[5 empty] | Executed | 206 | passed
Systems | passed | MASTER-18467 - For each memory location, the hearing aid shall allow E2E Directionality to be e... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 11 | passed
Systems | passed | MASTER-55396 - The WiCROS Transmitter shall use the following settings for Directionality: * En... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 148 | passed
Systems | passed | MASTER-18471 - When in Omnidirectional State, the hearing aid shall have the Omnidirectional Po... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 31 | passed
Systems | passed | MASTER-18472 - When in HF Directional State, the hearing aid shall have Omnidirectional Polar P... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 32 | passed
Systems | passed | MASTER-18473 - When in Full Directional State with Null Steering disabled, the hearing aid shal... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 33 | passed
Systems | passed | MASTER-18476 - Upon power-up, the hearing aid shall apply an initial directional state based on... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 39 | passed
[5 empty] | Executed | 46 | passed
[5 empty] | Executed | 53 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-18477 - Upon memory change, the hearing aid shall switch to an initial directional state... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 72 | passed
[5 empty] | Executed | 75 | passed
[5 empty] | Executed | 78 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 85 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 89 | passed
Systems | passed | MASTER-18481 - When E2E Directionality is enabled, the hearing aid shall provide an indication... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 20 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18482 - The hearing aid shall begin transition to Omnidirectional State within the Time... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 36 | passed
Systems | passed | MASTER-18483 - The hearing aid shall begin transition to HF Directional state within the Time C... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 20 | passed
[5 empty] | Executed | 31 | passed
Systems | passed | MASTER-18484 - The hearing aid shall begin transition to Full Directional state within the Time... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 29 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18486 - When a directional state transition is in progress, the hearing aid shall comple... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 20 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-18617 - When the polar pattern is not Omnidirectional, for each MECO channel, the hearin... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 95 | passed
[5 empty] | Executed | 111 | passed
Systems | passed | MASTER-18622 - When Noise Control is On, the hearing aid shall set its directionality state to... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 67 | passed
Systems | passed | MASTER-18638 - When the input level is less than Tk1, the hearing aid shall provide amplificati... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 18 | passed
Systems | passed | MASTER-18640 - When the input level is greater than or equal to Tk1 and less than Tk2, the hear... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 29 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-18642 - When the input level is greater than or equal to Tk2 and less than Tk3, the hear... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 29 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-18644 - When the input level is greater than or equal to Tk3, the hearing aid shall prov... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 29 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-18681 - A request for a volume control increase or decrease shall result in a correspond... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 20 | passed
Systems | passed | MASTER-18682 - A request for a volume control increase or decrease shall not affect the OCL com... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 76 | passed
Systems | passed | MASTER-18684 - When configured for a linear response, OCL is not engaged, adjacent MECO channel... | MASTER-34431 - AV111 - MECO | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 93 | passed
Systems | passed | MASTER-18704 - Upon entry into the Local Telephone Present state, when E2E ATR Memory Change is... | MASTER-28144 - SV025 - E2E User Control | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 12 | passed
|  |  | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 32 | passed
Systems | passed | MASTER-18702 - When the hearing aid detects a telephone, it shall enter the Local Telephone Pre... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 39 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-18701 - The hearing aid shall apply the Local Telephone Present, Opposite Ear Telephone... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-18705 - Upon transition from the Local Telephone Present state to the Telephone Not Pres... | MASTER-28144 - SV025 - E2E User Control | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
[5 empty] | Executed | 19 | passed
Systems | passed | MASTER-18706 - If the hearing aid is configured to use the Autophone feature, then upon entry i... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
Systems | passed | MASTER-18707 - If the hearing aid is configured to use the Autocoil feature, then upon entry in... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 66 | passed
Systems | passed | MASTER-18708 - When the Telephone indicator is enabled, on transition from the Telephone Not Pr... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 61 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 16 | passed
|  |  | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 102 | passed
[5 empty] | Executed | 110 | passed
Systems | passed | MASTER-18710 - If the patient does not change the active memory location while in the Local Tel... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
Systems | passed | MASTER-18711 - If the patient activates a different memory location while in the Local Telephon... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 45 | passed
Systems | passed | MASTER-18712 - On transition from the Opposite Ear Telephone Present state to the Local Telepho... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 49 | passed
Systems | passed | MASTER-18713 - When automatic telephone functionality has completed, an iOS audio stream was in... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 78 | passed
[5 empty] | Executed | 78 | passed
[5 empty] | Executed | 78 | passed
|  |  | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 47 | passed
Systems | passed | MASTER-18730 - When a telephone memory is manually activated and the Telephone indicator is ena... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 60 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 15 | passed
|  |  | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 79 | passed
[5 empty] | Executed | 88 | passed
[5 empty] | Executed | 132 | passed
Systems | passed | MASTER-18736 - The hearing aid shall allow Microphone/Telecoil (M/T) operation to be enabled an... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 59 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 117 | passed
Systems | passed | MASTER-18740 - For the Autocoil and Telecoil memory environments, the hearing aid shall allow t... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 60 | passed
[5 empty] | Executed | 73 | passed
[5 empty] | Executed | 127 | passed
Systems | passed | MASTER-18742 - When M/T operation is enabled, the hearing aid shall attenuate its microphone si... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 67 | passed
[5 empty] | Executed | 81 | passed
Systems | passed | MASTER-24800 - When M/T operation is enabled, the hearing aid shall operate in the Omnidirectio... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 61 | passed
[5 empty] | Executed | 74 | passed
Systems | passed | MASTER-50860 - The hearing aid shall allow data logging to be enabled and disabled. | MASTER-34494 - SV129 - Data Log | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 69 | passed
[5 empty] | Executed | 71 | passed
Systems | passed | MASTER-18749 - When data logging is enabled, the hearing aid shall provide data logging for the... | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 46 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 79 | passed
[5 empty] | Executed | 79 | passed
|  |  | MASTER-34494 - SV129 - Data Log | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 80 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 84 | passed
[5 empty] | Executed | 87 | passed
|  |  | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 54 | passed
Systems | passed | MASTER-18758 - When data logging is enabled, the hearing aid shall provide data logging while p... | MASTER-34462 - SV204 - Rechargeable HA Feature testing | IP9 - General IP9 Testing & Regression | Executed | 32 | passed
|  |  | MASTER-34494 - SV129 - Data Log | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 80 | passed
[5 empty] | Executed | 84 | passed
Systems | passed | MASTER-51353 - When the hearing aid receives a request to clear the manual alert data, the hear... | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 48 | passed
[5 empty] | Executed | 50 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 81 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 83 | passed
Systems | passed | MASTER-60743 - When enabled, the hearing aid shall present the audible indicators as defined in... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 71 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 73 | passed
[5 empty] | Executed | 74 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 78 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 58 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 75 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 83 | passed
[5 empty] | Executed | 89 | passed
[5 empty] | Executed | 94 | passed
[5 empty] | Executed | 97 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 156 | passed
[5 empty] | Executed | 170 | passed
[5 empty] | Executed | 173 | passed
[5 empty] | Executed | 175 | passed
[5 empty] | Executed | 177 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 15 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 48 | passed
[5 empty] | Executed | 52 | passed
|  |  | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 19 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 37 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 52 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 39 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 52 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 48 | passed
[5 empty] | Executed | 52 | passed
[5 empty] | Executed | 56 | passed
[5 empty] | Executed | 58 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 191 | passed
[5 empty] | Executed | 194 | passed
[5 empty] | Executed | 198 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 71 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 71 | passed
|  |  | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-60742 - When level dependence is enabled and the ambient sound level plus the Speech Ind... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 35 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 48 | passed
Systems | passed | MASTER-60745 - When an indicator is configured as a tone indicator, the hearing aid shall suppo... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 27 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 70 | passed
[5 empty] | Executed | 90 | passed
[5 empty] | Executed | 164 | passed
[5 empty] | Executed | 175 | passed
Systems | passed | MASTER-60748 - When an audible indicator is presented, the hearing aid shall only output the in... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 71 | passed
[5 empty] | Executed | 73 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 67 | passed
[5 empty] | Executed | 172 | passed
[5 empty] | Executed | 200 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 9 | passed
[5 empty] | Executed | 77 | passed
|  |  | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 19 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 38 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 46 | passed
[5 empty] | Executed | 51 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 47 | passed
[5 empty] | Executed | 51 | passed
[5 empty] | Executed | 55 | passed
[5 empty] | Executed | 59 | passed
[5 empty] | Executed | 64 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 71 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 71 | passed
|  |  | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-60749 - When the Cue indicator is enabled for a speech indicator, the hearing aid shall... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 52 | passed
[5 empty] | Executed | 55 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 92 | passed
Systems | passed | MASTER-19493 - The WiCROS Partner hearing aid shall present the indicators associated with the... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 12 | passed
Systems | passed | MASTER-60754 - When the Low Battery Warning, Low Battery Shutdown and Manual Power Off indicato... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 58 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-60751 - The hearing aid shall have a configurable base presentation level for each of th... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 141 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 54 | passed
[5 empty] | Executed | 56 | passed
Systems | passed | MASTER-60752 - The default base presentation level of speech indicators shall be equivalent to... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
|  |  | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 15 | passed
Systems | passed | MASTER-60756 - The default base presentation level of tone and tone series indicators shall be... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 2 | passed
Systems | passed | MASTER-60757 - Upon request from the fitting software, the hearing aid shall present the specif... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 54 | passed
[5 empty] | Executed | 56 | passed
Systems | passed | MASTER-60755 - When level dependence is not enabled, the hearing aid shall present the indicato... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 81 | passed
Systems | passed | MASTER-18964 - When a level dependence adjustment has been determined for an indicator, the hea... | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 25 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
Systems | passed | MASTER-18966 - Level dependence shall be disabled while the hearing aid is presenting a streame... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 72 | passed
Systems | passed | MASTER-60759 - When level dependence is enabled and the ambient sound level plus 5 dB is less t... | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 29 | passed
Systems | passed | MASTER-60760 - When level dependence is enabled and the ambient sound level plus 5 dB is greate... | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 30 | passed
Systems | passed | MASTER-60763 - When level dependence is enabled and the ambient sound level plus 5 dB is greate... | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 31 | passed
Systems | passed | MASTER-60765 - When level dependence is enabled and the ambient sound level plus 5 dB is less t... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 63 | passed
[5 empty] | Executed | 85 | passed
Systems | passed | MASTER-60766 - When level dependence is enabled and the ambient sound level plus 5 dB is greate... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 68 | passed
[5 empty] | Executed | 86 | passed
Systems | passed | MASTER-60771 - When ANCS indicators are enabled, upon receipt of a notification from an iOS dev... | MASTER-42132 - SV178 - Mobile Application - Apple Notification Center Service (ANCS) | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 24 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-60772 - Upon request from the mobile application, the hearing aid shall present the spec... | MASTER-42132 - SV178 - Mobile Application - Apple Notification Center Service (ANCS) | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 25 | passed
Systems | passed | MASTER-60770 - The manufacturing software shall disable ANCS indicators in the hearing aid. | MASTER-42132 - SV178 - Mobile Application - Apple Notification Center Service (ANCS) | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 16 | passed
Systems | passed | MASTER-19149 - The hearing aid shall inhibit power-on when the battery voltage is below 1.13 V. | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 13 | passed
Systems | passed | MASTER-19153 - The hearing aid shall allow a Power-on Delay Time to be set, ranging from 0 seco... | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 11 | passed
|  |  | MASTER-34467 - SV202 - Rechargeable HA functionality | Z22A - New Rechargeable Battery | Executed | 25 | passed
Systems | passed | MASTER-19155 - The hearing aid shall power-up with its audio muted. | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 23 | passed
[5 empty] | Executed | 29 | passed
Systems | passed | MASTER-19156 - After application of power and completion of initialization, the hearing aid sha... | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 18 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 30 | passed
Systems | passed | MASTER-19158 - The hearing aid shall complete initialization within 5 seconds of the applicatio... | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 17 | passed
Systems | passed | MASTER-19162 - When audio streaming is not active, and the hearing aid is not in a fitting sess... | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 38 | passed
[5 empty] | Executed | 55 | passed
[5 empty] | Executed | 57 | passed
Systems | passed | MASTER-19163 - When audio streaming is active, or the hearing aid is in a fitting session, and... | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-58687 - When the zinc-air battery voltage of the WiCROS Transmitter falls below 1.13 vol... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 60 | passed
Systems | passed | MASTER-19164 - Thirty minutes after the streaming or non-streaming low battery criteria have be... | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 28 | passed
[5 empty] | Executed | 46 | passed
[5 empty] | Executed | 62 | passed
Systems | passed | MASTER-53067 - Thirty minutes after the WiCROS Transmitter low battery criteria have been met,... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 62 | passed
Systems | passed | MASTER-53068 - Thirty minutes after the low battery criteria have been met, the WiCROS Transmit... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 62 | passed
Systems | passed | MASTER-19179 - The hearing aid shall operate as specified herein until the point of powering do... | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 29 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-19181 - Upon command from the fitting software, the hearing aid shall transition to a me... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 12 | passed
Systems | passed | MASTER-51113 - Upon entry into a wireless fitting session, the hearing aid shall configure its... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 164 | passed
[5 empty] | Executed | 167 | passed
[5 empty] | Executed | 168 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 170 | passed
[5 empty] | Executed | 171 | passed
[5 empty] | Executed | 172 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 209 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 213 | passed
[5 empty] | Executed | 214 | passed
[5 empty] | Executed | 215 | passed
[5 empty] | Executed | 216 | passed
[5 empty] | Executed | 217 | passed
[5 empty] | Executed | 232 | passed
Systems | passed | MASTER-51109 - The hearing aid shall support E2E Phone Streaming while in a wireless fitting se... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 13 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-51110 - The WiCROS system shall support WiCROS audio streaming while in a wireless fitti... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 28 | passed
Systems | passed | MASTER-19204 - The hearing aid shall have a Wireless Fitting Session Timeout of 20 seconds. Not... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 178 | passed
Systems | passed | MASTER-19205 - When a hearing aid in a wireless fitting session receives no communication from... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 183 | passed
[5 empty] | Executed | 186 | passed
[5 empty] | Executed | 188 | passed
[5 empty] | Executed | 189 | passed
[5 empty] | Executed | 191 | passed
[5 empty] | Executed | 193 | passed
[5 empty] | Executed | 194 | passed
Systems | passed | MASTER-19208 - Upon receipt of a command from the fitting software to re-initialize, the hearin... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 183 | passed
[5 empty] | Executed | 186 | passed
[5 empty] | Executed | 188 | passed
[5 empty] | Executed | 189 | passed
[5 empty] | Executed | 191 | passed
[5 empty] | Executed | 193 | passed
[5 empty] | Executed | 194 | passed
Systems | passed | MASTER-19216 - The hearing aid shall allow the fitting software to perform a field firmware upd... | MASTER-34468 - SV069 - Verify FFU | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 42 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 71 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 88 | passed
Systems | passed | MASTER-19217 - The hearing aid shall allow the fitting software to perform a field firmware upd... | MASTER-34468 - SV069 - Verify FFU | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 90 | passed
Systems | passed | MASTER-42144 - The hearing aid shall allow the mobile application to perform a field firmware u... | MASTER-47171 - SV227 - Mobile FFU | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
Systems | passed | MASTER-42145 - In a binaural fitting session, a field firmware update of both hearing aids shal... | MASTER-34468 - SV069 - Verify FFU | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 45 | passed
Systems | passed | MASTER-19593 - The hearing aid shall turn Noise Control On and sound the Noise Control On indic... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 82 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 106 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 124 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-19594 - The hearing aid shall turn Noise Control Off and sound the Noise Control Off ind... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 85 | passed
[5 empty] | Executed | 89 | passed
[5 empty] | Executed | 108 | passed
Systems | passed | MASTER-19596 - When Noise Control is On, the hearing aid shall turn Noise Control Off when any... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 40 | passed
[5 empty] | Executed | 124 | passed
[5 empty] | Executed | 128 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 133 | passed
[5 empty] | Executed | 140 | passed
[5 empty] | Executed | 141 | passed
Systems | passed | MASTER-19598 - Upon power-up, the hearing aid shall turn Noise Control Off. | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
Systems | passed | MASTER-19599 - Muting and unmuting the hearing aid via a User Control Mute Request shall have n... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 46 | passed
[5 empty] | Executed | 48 | passed
[5 empty] | Executed | 51 | passed
[5 empty] | Executed | 53 | passed
[5 empty] | Executed | 55 | passed
[5 empty] | Executed | 57 | passed
Systems | passed | MASTER-19602 - When Noise Control is On, the hearing aid shall switch to Full Directional state... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 67 | passed
Systems | passed | UTICA-6854 - When Noise Control is On, the hearing aid shall apply MNR noise reduction in acc... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 82 | passed
[5 empty] | Executed | 87 | passed
Systems | passed | MASTER-19609 - For each memory location, the hearing aid shall allow Frequency Translation to b... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 26 | passed
Systems | passed | MASTER-19614 - For each memory location, the hearing aid shall allow the Frequency Translation... | MASTER-26425 - SV121A - Mfg Suite | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 31 | passed
Systems | passed | MASTER-19615 - When Frequency Translation is enabled, the hearing aid shall only perform Freque... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 69 | passed
Systems | passed | MASTER-19617 - When Frequency Translation is enabled, the hearing aid shall output no audio in... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 67 | passed
Systems | passed | MASTER-19619 - When Frequency Translation is enabled, the hearing aid shall move a translated p... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 63 | passed
[5 empty] | Executed | 65 | passed
Systems | passed | MASTER-19620 - When Frequency Translation is enabled, the hearing aid shall increase the gain o... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 66 | passed
Systems | passed | MASTER-19621 - The hearing aid shall not process tone indicators through the Frequency Translat... | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 42 | passed
Systems | passed | MASTER-19636 - For each memory location, the hearing aid shall allow Auto Acclimatization to be... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 13 | passed
Systems | passed | MASTER-19638 - For each MECO channel, the hearing aid shall allow a 6-step gain adaptation plan... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 66 | passed
Systems | passed | MASTER-19641 - When Auto Acclimatization is enabled, the hearing aid shall initiate the feature... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 44 | passed
Systems | passed | MASTER-19642 - When Auto Acclimatization is enabled, the hearing aid shall initiate the feature... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 48 | passed
Systems | passed | MASTER-19643 - When the time spent on the current step of the Auto Acclimatization adaptation p... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 66 | passed
Systems | passed | MASTER-19645 - The hearing aid shall allow the fitting software to pause and resume Auto Acclim... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 77 | passed
Systems | passed | MASTER-19646 - When Auto Acclimatization is paused, the hearing aid shall suspend the adaptatio... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 86 | passed
Systems | passed | MASTER-19647 - When Auto Acclimatization is resumed, the hearing aid shall continue the adaptat... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 88 | passed
Systems | passed | MASTER-19649 - When the Auto Acclimatization adaptation plan is in progress and a fitting sessi... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 75 | passed
Systems | passed | MASTER-19655 - The hearing aid shall allow a value to be set indicating that the Resistive ID f... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 72 | passed
[5 empty] | Executed | 78 | passed
Systems | passed | MASTER-19657 - The hearing aid shall be capable of uniquely identifying the following resistor... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 87 | passed
Systems | passed | MASTER-19662 - The hearing aid shall inhibit the RID function when the magnetic switch is activ... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 58 | passed
Systems | passed | MASTER-19664 - The hearing aid shall provide current RID values to the fitting software upon co... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 34 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-56289 - Upon command from the fitting software when a RID-capable receiver cable is conn... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 14 | passed
[5 empty] | Executed | 78 | passed
Systems | passed | MASTER-19666 - During a fitting session, when there is no receiver cable connected to the heari... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 14 | passed
Systems | passed | MASTER-19668 - During a fitting session, when an RID-capable receiver cable is connected to the... | MASTER-26718 - SV086 - Resistive ID | IP4 - General IP4 Testing (Utica Features & Tulum Bug Fixes) | Executed | 20 | passed
Systems | passed | MASTER-51530 - The hearing aid shall allow pairing with the following number of devices listed... | MASTER-50731 - SV238 - Multi-Accessory Interactions | IP6 - General IP6 Testing | Executed | 108 | passed
[5 empty] | Executed | 134 | passed
[5 empty] | Executed | 135 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 108 | passed
[5 empty] | Executed | 134 | passed
[5 empty] | Executed | 135 | passed
[5 empty] | Executed | 136 | passed
Systems | passed | MASTER-51531 - If the hearing aid is already paired with the maximum supported number of device... | MASTER-50731 - SV238 - Multi-Accessory Interactions | IP6 - General IP6 Testing | Executed | 130 | passed
[5 empty] | Executed | 132 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 132 | passed
Systems | passed | MASTER-19234 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 15 | passed
|  |  | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 29 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 69 | passed
Systems | passed | MASTER-19235 - Upon receipt of a pairing request from a BLE Central within 3 minutes of power a... | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 24 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-19236 - Upon receipt of a pairing request from a BLE Central greater than 3 minutes afte... | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 21 | passed
Systems | passed | MASTER-19237 - When successfully paired with a BLE Central, the hearing aid shall communicate w... | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 19 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-19239 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 15 | passed
Systems | passed | MASTER-19240 - Upon receipt of a pairing request from a 2.4 GHz Remote Control within 3 minutes... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-19241 - Upon receipt of a pairing request from a 2.4 GHz Remote Control greater than 3 m... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 39 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-50664 - Upon receipt of a pairing request from a 2.4 GHz Remote Control within 30 second... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 89 | passed
|  |  | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 14 | passed
Systems | passed | MASTER-50665 - Upon receipt of a pairing request from a 2.4 GHz Remote Control greater than 30... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 86 | passed
Systems | passed | MASTER-19242 - When successfully paired with a 2.4 GHz Remote Control, the hearing aid shall co... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 48 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 74 | passed
[5 empty] | Executed | 75 | passed
Systems | passed | MASTER-25416 - When successfully paired with a 2.4 GHz Remote Control, the hearing aid shall pr... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 45 | passed
Systems | passed | MASTER-42151 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-42152 - Upon receipt of a pairing request from a 2.4 GHz TV Streamer within 3 minutes of... | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-42153 - Upon receipt of a pairing request from a 2.4 GHz TV Streamer greater than 3 minu... | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-50666 - Upon receipt of a pairing request from a 2.4 GHz TV Streamer within 30 seconds o... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 11 | passed
|  |  | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-50667 - Upon receipt of a pairing request from a 2.4 GHz TV Streamer greater than 30 sec... |  |  | Executed | N/A | passed
Systems | passed | MASTER-42154 - When successfully paired with a 2.4 GHz TV Streamer, the hearing aid shall commu... | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-42156 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 58 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-42157 - Upon receipt of a pairing request from a 2.4 GHz MFA within 3 minutes of power a... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 33 | passed
[5 empty] | Executed | 78 | passed
Systems | passed | MASTER-42158 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 3 minutes afte... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 29 | passed
Systems | passed | MASTER-50668 - Upon receipt of a pairing request from a 2.4 GHz MFA within 30 seconds of of a p... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 8 | passed
|  |  | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 93 | passed
[5 empty] | Executed | 115 | passed
Systems | passed | MASTER-50669 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 30 seconds aft... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 90 | passed
[5 empty] | Executed | 112 | passed
Systems | passed | MASTER-42159 - When successfully paired with a 2.4 GHz MFA, the hearing aid shall communicate w... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 43 | passed
[5 empty] | Executed | 55 | passed
[5 empty] | Executed | 64 | passed
Systems | passed | MASTER-50863 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 46 | passed
[5 empty] | Executed | 48 | passed
Systems | passed | MASTER-50864 - Upon receipt of a pairing request from a 2.4 GHz Remote Mic within 3 minutes of... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-50861 - Upon receipt of a pairing request from a 2.4 GHz Remote Mic greater than 3 minut... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 22 | passed
Systems | passed | MASTER-50862 - Upon receipt of a pairing request from a 2.4 GHz Remote Mic within 60 seconds of... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 13 | passed
Systems | passed | MASTER-50865 - Upon receipt of a pairing request from a 2.4 GHz Remote Mic greater than 60 seco... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 11 | passed
Systems | passed | MASTER-50866 - When successfully paired with a 2.4 GHz Remote Mic, the hearing aid shall commun... | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 34 | passed
Systems | passed | MASTER-62845 - The hearing aid shall take no action that would result in a change in its hearin... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 132 | passed
Systems | passed | MASTER-62846 - Upon receipt of a pairing request from a 2.4 GHz Microphone Pod within 3 minutes... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 17 | passed
Systems | passed | MASTER-62843 - Upon receipt of a pairing request from a 2.4 GHz Microphone Pod greater than 3 m... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
Systems | passed | MASTER-62844 - Upon receipt of a pairing request from a 2.4 GHz Microphone Pod within 30 second... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
Systems | passed | MASTER-62847 - Upon receipt of a pairing request from a 2.4 GHz Microphone Pod greater than 30... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
Systems | passed | MASTER-62848 - When successfully paired with a 2.4 GHz Microphone Pod, the hearing aid shall co... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-50670 - Upon receipt of a pairing deletion command for one of the following accessory de... | MASTER-46421 - SV224 - 2.4 Ghz TV streamer | Model S IP 6-5 | Executed | N/A | passed
Systems | passed | MASTER-19249 - The hearing aid shall allow another wireless hearing aid to be assigned as its E... | MASTER-47045 - AV025 - Automated E2E User Control | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 48 | passed
Systems | passed | MASTER-50036 - For each memory location, the hearing aid shall allow one of the following optio... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 17 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 58 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 70 | passed
Systems | passed | MASTER-58670 - The hearing aid shall allow one of the following options to be set for Initial A... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 127 | passed
Systems | passed | MASTER-63411 - For audio streams from iOS and Android Devices, the hearing aid shall allow Volu... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 176 | passed
[5 empty] | Executed | 185 | passed
[5 empty] | Executed | 195 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 185 | passed
[5 empty] | Executed | 195 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 185 | passed
[5 empty] | Executed | 195 | passed
Systems | passed | MASTER-42160 - The hearing aid shall process audio streams in the following priority: Stream Ty... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 94 | passed
[5 empty] | Executed | 99 | passed
[5 empty] | Executed | 100 | passed
[5 empty] | Executed | 102 | passed
[5 empty] | Executed | 103 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 122 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 133 | passed
[5 empty] | Executed | 134 | passed
[5 empty] | Executed | 137 | passed
[5 empty] | Executed | 154 | passed
[5 empty] | Executed | 156 | passed
[5 empty] | Executed | 161 | passed
[5 empty] | Executed | 162 | passed
[5 empty] | Executed | 165 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 171 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 178 | passed
[5 empty] | Executed | 181 | passed
[5 empty] | Executed | 184 | passed
[5 empty] | Executed | 188 | passed
[5 empty] | Executed | 190 | passed
Systems | passed | MASTER-19267 - When an audio stream is in progress and notification of a higher priority audio... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 94 | passed
[5 empty] | Executed | 95 | passed
[5 empty] | Executed | 99 | passed
[5 empty] | Executed | 100 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 122 | passed
[5 empty] | Executed | 128 | passed
[5 empty] | Executed | 133 | passed
[5 empty] | Executed | 154 | passed
[5 empty] | Executed | 161 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 174 | passed
[5 empty] | Executed | 178 | passed
[5 empty] | Executed | 184 | passed
[5 empty] | Executed | 190 | passed
[5 empty] | Executed | 210 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 214 | passed
[5 empty] | Executed | 216 | passed
Systems | passed | MASTER-19268 - When an audio stream is in progress and notification of a lower priority audio s... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 99 | passed
[5 empty] | Executed | 165 | passed
[5 empty] | Executed | 171 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 181 | passed
[5 empty] | Executed | 188 | passed
[5 empty] | Executed | 200 | passed
[5 empty] | Executed | 201 | passed
[5 empty] | Executed | 202 | passed
[5 empty] | Executed | 203 | passed
[5 empty] | Executed | 205 | passed
[5 empty] | Executed | 207 | passed
[5 empty] | Executed | 209 | passed
Systems | passed | MASTER-60080 - When an audio stream is in progress and notification of a lower priority audio s... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 165 | passed
[5 empty] | Executed | 171 | passed
[5 empty] | Executed | 176 | passed
[5 empty] | Executed | 181 | passed
[5 empty] | Executed | 188 | passed
[5 empty] | Executed | 203 | passed
[5 empty] | Executed | 207 | passed
[5 empty] | Executed | 209 | passed
Systems | passed | MASTER-59912 - When an audio stream is being received from an iOS Device, and the iOS Device re... | MASTER-73851 - SV284 - LEA2 Mic Streaming Test Case | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 27 | passed
[5 empty] | Executed | 29 | passed
Systems | passed | MASTER-59915 - The hearing aid shall stop transmitting the microphone stream upon request from... | MASTER-73851 - SV284 - LEA2 Mic Streaming Test Case | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 44 | passed
Systems | passed | MASTER-59918 - When transmitting a microphone stream to an iOS Device, and the corresponding in... | MASTER-73851 - SV284 - LEA2 Mic Streaming Test Case | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 26 | passed
Systems | passed | MASTER-19252 - When notification of an audio stream is received from an iOS Device, Android Dev... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 58 | passed
[5 empty] | Executed | 58 | passed
[5 empty] | Executed | 58 | passed
|  |  | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 28 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 43 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 47 | passed
Systems | passed | MASTER-19253 - When notification of an audio stream is received from an iOS Device, Android Dev... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 81 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 81 | passed
Systems | passed | MASTER-42167 - When an audio stream is being received from an iOS Device, Android Device, or an... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 95 | passed
[5 empty] | Executed | 100 | passed
[5 empty] | Executed | 219 | passed
[5 empty] | Executed | 222 | passed
[5 empty] | Executed | 226 | passed
[5 empty] | Executed | 230 | passed
Systems | passed | MASTER-42168 - When an audio stream from an iOS Device or Android Device has been interrupted d... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 102 | passed
[5 empty] | Executed | 221 | passed
Systems | passed | MASTER-42169 - When an audio stream from an iOS Device or Android Device has been interrupted d... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 125 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 224 | passed
Systems | passed | MASTER-50296 - When an audio stream from an accessory device has been interrupted due to the he... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 97 | passed
[5 empty] | Executed | 228 | passed
[5 empty] | Executed | 232 | passed
Systems | passed | MASTER-19278 - When an audio stream is not being received from an accessory device, the hearing... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 28 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 45 | passed
[5 empty] | Executed | 48 | passed
Systems | passed | MASTER-42178 - When a Start Accessory Stream Request has been received and an audio stream from... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 30 | passed
Systems | passed | MASTER-42180 - When a Start Accessory Stream Request has been received and an audio stream from... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 61 | passed
Systems | passed | MASTER-42182 - When a Stop Accessory Stream Request has been received, the hearing aid shall: *... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 73 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 361 | passed
[5 empty] | Executed | 369 | passed
Systems | passed | MASTER-42187 - Upon receipt of a request to initiate audio streaming from a specific paired acc... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 30 | passed
Systems | passed | MASTER-42188 - When a request to initiate audio streaming from a specific paired accessory has... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 32 | passed
[5 empty] | Executed | 36 | passed
Systems | passed | MASTER-42189 - When a request to initiate audio streaming from a specific paired accessory has... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 30 | passed
Systems | passed | MASTER-42192 - Upon receipt of a request to terminate audio streaming from a specific paired ac... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 45 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 361 | passed
[5 empty] | Executed | 369 | passed
Systems | passed | MASTER-42194 - When audio streaming from an iOS Device, Android Device, or an accessory device... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 33 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 361 | passed
[5 empty] | Executed | 369 | passed
Systems | passed | MASTER-19308 - The hearing aid shall allow Auto-Streaming to be enabled and disabled. Note: Whi... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 44 | passed
[5 empty] | Executed | 85 | passed
[5 empty] | Executed | 99 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 85 | passed
[5 empty] | Executed | 99 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 85 | passed
[5 empty] | Executed | 99 | passed
Systems | passed | MASTER-58671 - For the microphone accessory stream input type of the 2.4 GHz Microphone Pod acc... | MASTER-61192 - SV258 - Nuance Test Procedure New | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 129 | passed
Systems | passed | MASTER-19314 - The hearing aid shall allow the Auto-Streaming memory to be included in rotation... | MASTER-26374 - SV002B - User Control - Memory | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 135 | passed
[5 empty] | Executed | 138 | passed
Systems | passed | MASTER-19318 - When Auto-Streaming is enabled, an audio stream is received and the hearing aid... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 96 | passed
[5 empty] | Executed | 96 | passed
[5 empty] | Executed | 96 | passed
Systems | passed | MASTER-19321 - When a memory change request is received as a result of a user control input or... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 64 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 64 | passed
Systems | passed | MASTER-19322 - When a memory change request is received as a result of automatic telephone func... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 66 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-19326 - When an audio stream ends and a memory change has not occurred during the stream... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 61 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-19327 - When an audio stream ends and a memory change occurred during the stream as a re... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 64 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 64 | passed
Systems | passed | MASTER-19328 - When an audio stream ends, a memory change occurred during the stream as a resul... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 68 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 68 | passed
Systems | passed | MASTER-19329 - When an audio stream ends, a memory change occurred during the stream as a resul... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 73 | passed
[5 empty] | Executed | 73 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-19352 - For each memory location, the hearing aid shall allow Initiate E2E Phone Streami... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 8 | passed
[5 empty] | Executed | 14 | passed
[5 empty] | Executed | 44 | passed
Systems | passed | MASTER-19353 - When an Autocoil or Telecoil memory is activated and Initiate E2E Phone Streamin... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 71 | passed
Systems | passed | MASTER-19354 - When an Autophone or Telephone memory is activated and Initiate E2E Phone Stream... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 57 | passed
Systems | passed | MASTER-19355 - The hearing aid shall include its current memory location in the message initiat... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 70 | passed
Systems | passed | MASTER-19357 - For each memory location, the hearing aid shall allow one of the following Recei... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 18 | passed
Systems | passed | MASTER-19358 - For each memory location, the hearing aid shall have an E2E Phone Streaming Micr... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 84 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 95 | passed
Systems | passed | MASTER-19360 - When the current memory location is configured as E2E Phone Streaming Only, the... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 58 | passed
Systems | passed | MASTER-19362 - When the current memory location is configured as E2E Phone Streaming + Micropho... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 71 | passed
[5 empty] | Executed | 84 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 95 | passed
Systems | passed | MASTER-19364 - When notification of an E2E Phone audio stream is received from the E2E Partner... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 188 | passed
[5 empty] | Executed | 190 | passed
Systems | passed | MASTER-19365 - The hearing aid shall begin presentation of the received E2E Phone audio stream... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 57 | passed
[5 empty] | Executed | 71 | passed
Systems | passed | MASTER-19366 - When the hearing aid receiving an E2E Phone audio stream transitioned to the cur... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 76 | passed
Systems | passed | MASTER-19368 - When notification of an E2E Phone audio stream is received from the E2E Partner... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 102 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-19370 - When E2E Phone Streaming is in progress, the presentation of the audio signal th... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 59 | passed
[5 empty] | Executed | 72 | passed
Systems | passed | MASTER-19371 - The E2E Phone audio stream shall have a frequency range of at least 400 Hz to 34... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 150 | passed
[5 empty] | Executed | 161 | passed
Systems | passed | MASTER-42199 - The hearing aid shall respond to a user command, regardless of input method, whi... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 122 | passed
[5 empty] | Executed | 124 | passed
[5 empty] | Executed | 126 | passed
Systems | passed | MASTER-51646 - The hearing aid shall support E2E Phone Streaming while in a wireless fitting se... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 13 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-19383 - For each memory location, the WiCROS system shall allow one of the following opt... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 27 | passed
Systems | passed | MASTER-19393 - Prior to presentation of the WiCROS audio stream, the WiCROS Partner hearing aid... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 12 | passed
Systems | passed | MASTER-19388 - Upon transition from a memory location in which WiCROS streaming is disabled to... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 26 | passed
Systems | passed | MASTER-19394 - When the current memory location is configured as CROS, the WiCROS Partner heari... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 28 | passed
Systems | passed | MASTER-19395 - When the current memory location is configured as BiCROS, the WiCROS Partner hea... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
Systems | passed | MASTER-51115 - When the current memory location is configured as Hearing Aid Only, the WiCROS P... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
Systems | passed | MASTER-19397 - When a WiCROS audio stream is being presented in a memory location in which WiCR... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 30 | passed
Systems | passed | MASTER-19399 - When presenting a mix of the WiCROS Transmitter streamed audio signal and the Wi... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 81 | passed
Systems | passed | MASTER-19405 - The WiCROS system shall only respond to user control requests to adjust the WiCR... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 83 | passed
[5 empty] | Executed | 88 | passed
[5 empty] | Executed | 89 | passed
Systems | passed | MASTER-19408 - The WiCROS system shall allow the default Balance Control Level to be set. | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
|  |  | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 79 | passed
[5 empty] | Executed | 81 | passed
Systems | passed | MASTER-19409 - Upon power-up, the WiCROS system shall set the Balance Control Level to the prog... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 73 | passed
|  |  | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 81 | passed
Systems | passed | MASTER-19411 - Upon receipt of a request via user control to increase its WiCROS balance and th... |  |  | Executed | 120 | passed
Systems | passed | MASTER-19413 - Upon receipt of a request via user control to decrease its WiCROS balance and th... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 123 | passed
Systems | passed | MASTER-19417 - Upon a change in memory location, when Volume Reset on Memory Change is enabled,... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 109 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 133 | passed
[5 empty] | Executed | 135 | passed
[5 empty] | Executed | 158 | passed
[5 empty] | Executed | 160 | passed
Systems | passed | MASTER-19418 - Upon a change in memory location, when Volume Reset on Memory Change is disabled... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 114 | passed
|  |  | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 144 | passed
[5 empty] | Executed | 146 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 171 | passed
Systems | passed | MASTER-19420 - When transitioning from one WiCROS balance level to another as a result of user... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 84 | passed
Systems | passed | MASTER-19422 - The WiCROS system shall respond to a user command, regardless of input method, w... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 16 | passed
Systems | passed | MASTER-19424 - The WiCROS Transmitter shall preclude the presentation of any audio through the... | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 4 | passed
Systems | passed | MASTER-19539 - Upon completion of power on, the hearing aid shall provide status to the BLE Cen... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 105 | passed
[5 empty] | Executed | 111 | passed
[5 empty] | Executed | 105 | passed
[5 empty] | Executed | 111 | passed
[5 empty] | Executed | 105 | passed
[5 empty] | Executed | 111 | passed
|  |  | MASTER-42121 - SV170 - Mobile Application - Native Application and Controls | IP5 - General IP5 Testing | Executed | 35 | passed
[5 empty] | Executed | 40 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 53 | passed
[5 empty] | Executed | 59 | passed
Systems | passed | MASTER-19540 - Upon completion of a microphone volume change, the hearing aid shall provide sta... | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 19 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 19 | passed
[5 empty] | Executed | 26 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 30 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-19541 - Upon completion of a user control action resulting in the muting or unmuting of... | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 24 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 36 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 41 | passed
Systems | passed | MASTER-19542 - The hearing aid shall allow the Starkey Mobile App to read its Use Gain level. N... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 106 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 106 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 106 | passed
[5 empty] | Executed | 112 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 54 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-50084 - Upon completion of power on, the hearing aid shall provide status to the BLE Cen... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 27 | passed
Systems | passed | MASTER-50082 - Upon completion of a microphone volume change, the hearing aid shall provide sta... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 39 | passed
Systems | passed | MASTER-19544 - The hearing aid shall allow a value to be set indicating whether or not a memory... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
|  |  | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 77 | passed
[5 empty] | Executed | 88 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 76 | passed
[5 empty] | Executed | 90 | passed
[5 empty] | Executed | 76 | passed
[5 empty] | Executed | 90 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
[5 empty] | Executed | 34 | passed
Systems | passed | MASTER-19545 - Upon completion of power on, the hearing aid shall indicate to the BLE Central w... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 41 | passed
|  |  | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 77 | passed
[5 empty] | Executed | 88 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 76 | passed
[5 empty] | Executed | 90 | passed
[5 empty] | Executed | 76 | passed
[5 empty] | Executed | 90 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 20 | passed
[5 empty] | Executed | 34 | passed
Systems | passed | MASTER-19546 - Upon completion of a memory change, where the memory change request did not orig... | MASTER-42124 - SV172 - Mobile Application - Mobile Application and Controls | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 50 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 50 | passed
|  |  | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 43 | passed
Systems | passed | MASTER-19549 - Upon completion of power on, the hearing aid shall provide status to the BLE Cen... | MASTER-42127 - SV173 - Mobile Application - User Control | IP5 - General IP5 Testing | Executed | 30 | passed
[5 empty] | Executed | 30 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 43 | passed
[5 empty] | Executed | 43 | passed
Systems | passed | MASTER-19550 - Upon completion of a Tinnitus Therapy Stimulus Level change, the hearing aid sha... | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 114 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
[5 empty] | Executed | 114 | passed
[5 empty] | Executed | 116 | passed
[5 empty] | Executed | 118 | passed
[5 empty] | Executed | 120 | passed
Systems | passed | MASTER-19551 - Upon completion of a user control action resulting in the muting or unmuting of... | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 102 | passed
[5 empty] | Executed | 104 | passed
[5 empty] | Executed | 102 | passed
[5 empty] | Executed | 104 | passed
Systems | passed | MASTER-19553 - Upon completion of power on, the hearing aid shall provide battery status to the... | MASTER-42131 - SV177 - Mobile Application - BLE Central Battery Capacity Status | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 28 | passed
Systems | passed | MASTER-19554 - When audio streaming is active, the hearing aid shall provide battery status to... | MASTER-42131 - SV177 - Mobile Application - BLE Central Battery Capacity Status | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 61 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 67 | passed
[5 empty] | Executed | 70 | passed
Systems | passed | MASTER-19555 - When audio streaming is not active, the hearing aid shall provide battery status... | MASTER-42131 - SV177 - Mobile Application - BLE Central Battery Capacity Status | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 33 | passed
[5 empty] | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 42 | passed
Systems | passed | MASTER-51147 - The WiCROS system shall provide battery status to the BLE Central for both the W... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 105 | passed
Systems | passed | MASTER-19556 - When the criteria for a given battery status level have been met, that status le... | MASTER-42131 - SV177 - Mobile Application - BLE Central Battery Capacity Status | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 45 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-19558 - The hearing aid shall provide the enabled/disabled state of the Auto-Streaming f... | MASTER-42128 - SV174 - Mobile Application - Streaming and Stream Quality | Continuation 1 Mobile - Testing of Thrive Release for Utica Continuation 1 | Executed | 44 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 78 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 92 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 78 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 92 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 78 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 92 | passed
Systems | passed | MASTER-42201 - When streaming from an accessory device other than the OrCam Visual Assistant, t... | MASTER-51222 - SV166- Mobile Application - Accessories Screen | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 21 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 23 | passed
Systems | passed | MASTER-19560 - The hearing aid shall provide the version of its embedded firmware to the BLE Ce... | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 27 | passed
Systems | passed | MASTER-19562 - The hearing aid shall provide a status to the BLE Central when Noise Control is... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 40 | passed
[5 empty] | Executed | 43 | passed
Systems | passed | MASTER-50870 - Upon completion of power on, the WiCROS system shall provide status to the BLE C... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 81 | passed
Systems | passed | MASTER-50871 - Upon completion of a WiCROS Balance change, the WiCROS system shall provide stat... | MASTER-53146 - SV243 - NFMI WiCROS - Remote Control & Mobile App | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 84 | passed
Systems | passed | MASTER-50193 - When the BLE Central requests Brain Score data, the hearing aid shall transmit t... | MASTER-46418 - SV219 - Activity Tracker | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 31 | passed
Systems | passed | MASTER-60918 - The hearing aid shall provide the active/deactivated state of the On-Demand Adap... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 75 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 75 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-19564 - When the connection to a BLE Central has been lost, the hearing aid shall: * Rem... | MASTER-42123 - SV171A - Mobile Application - Pairing and Reconnections | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 47 | passed
Systems | passed | MASTER-62898 - The hearing aid shall detect the user's steps while walking with an accuracy of... | MASTER-46418 - SV219 - Activity Tracker | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 45 | passed
Systems | passed | MASTER-47583 - The hearing aid shall log the most recent 8 hours' worth of activity data. | MASTER-46418 - SV219 - Activity Tracker | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 117 | passed
[5 empty] | Executed | 118 | passed
Systems | passed | MASTER-51368 - When an E2E transmit alert command is received, the hearing aid shall: * Transmi... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 70 | passed
[5 empty] | Executed | 90 | passed
[5 empty] | Executed | 70 | passed
[5 empty] | Executed | 90 | passed
Systems | passed | MASTER-51411 - If the hearing aid receives an alert delivered receipt from BLE Central, or afte... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 93 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-51426 - If the hearing aid is presenting an accessory audio stream when sending an alert... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 65 | passed
[5 empty] | Executed | 65 | passed
Systems | passed | MASTER-51423 - If the hearing aid is presenting an audio stream from BLE Central and when the h... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 60 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-56883 - When connected to a BLE Central, upon receipt of a request to initiate the Perso... | MASTER-56939 - SV250 - Personal Voice Assistant | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 24 | passed
[5 empty] | Executed | 62 | passed
Systems | passed | MASTER-56885 - When not connected to a BLE Central, upon receipt of a request to initiate the P... | MASTER-56939 - SV250 - Personal Voice Assistant | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 44 | passed
[5 empty] | Executed | 78 | passed
Systems | passed | MASTER-56886 - Upon receipt of a Personal Voice Assistant activated response from the BLE Centr... | MASTER-56939 - SV250 - Personal Voice Assistant | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 20 | passed
[5 empty] | Executed | 58 | passed
Systems | passed | MASTER-56887 - Upon receipt of a Personal Voice Assistant failure response from the BLE Central... | MASTER-56939 - SV250 - Personal Voice Assistant | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 50 | passed
[5 empty] | Executed | 84 | passed
Systems | passed | MASTER-56888 - When a Personal Voice Assistant response is not received from the BLE Central in... | MASTER-56939 - SV250 - Personal Voice Assistant | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 47 | passed
[5 empty] | Executed | 81 | passed
Systems | passed | MASTER-57695 - When the following conditions have been met: * Auto Sleep is enabled * The heari... | MASTER-58171 - SV253- Auto Wake_sleep | IP6 - General IP6 Testing | Executed | 13 | passed
[5 empty] | Executed | 29 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-57696 - When utilizing Auto Sleep instead of nightly powering off the hearing aid, the h... | MASTER-58171 - SV253- Auto Wake_sleep | IP6 - General IP6 Testing | Executed | 78 | passed
Systems | passed | MASTER-60906 - Upon receipt of a request to activate On-Demand Adaptive Tuning, the hearing aid... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-60911 - Upon receipt of a request to activate On-Demand Adaptive Tuning, the hearing aid... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
Systems | passed | MASTER-61645 - In a binaural fitting, On-Demand Adaptive Tuning shall only be activated when th... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
Systems | passed | MASTER-61644 - In a binaural fitting, when the two hearing aids have classified the acoustic in... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 66 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | UTICA-25382 - When the acoustic input has been classified and On-Demand Adaptive Tuning remain... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 60 | passed
[5 empty] | Executed | 60 | passed
Systems | passed | MASTER-64595 - On request from the mobile application, the hearing aid shall return information... | MASTER-66967 - SV277C - DNN-Based Speech Enhancement on Phone - Datalogging | IP5 Mobile - IP5 Thrive Testing (New Utica Features) | Executed | 19 | passed
[5 empty] | Executed | 19 | passed
Systems | passed | MASTER-20295 - The hearing aid shall allow Magnetic Switch Memory Cycle to be enabled and disab... | MASTER-66945 - SV121B - Manufacturing Suite (For 2.4 Devices Only) | Amplifon - IP2 testing | Executed | 28 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 33 | passed
Systems | passed | MASTER-20296 - When Magnetic Switch Memory Cycle is enabled, the hearing aid shall recognize th... | MASTER-66945 - SV121B - Manufacturing Suite (For 2.4 Devices Only) | Amplifon - IP2 testing | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
ElectroAcoustics | passed | MASTER-19372 - For each third-octave band from 500 through 3150 Hz, the E2E Phone audio stream... | UTICA-23148 - Streaming Tests for Paris HA | EA_Paris_01 - Utica Features on Paris HA_EA DVT | Executed | 1 | passed
|  |  | UTICA-23161 - Streaming Tests for Quebec HA | EA_Quebec_01 - Utica Features on Quebec HA | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19400 - For each third-octave band from 500 through 3150 Hz, the WiCROS audio stream, as... | UTICA-23148 - Streaming Tests for Paris HA | EA_Paris_01 - Utica Features on Paris HA_EA DVT | Executed | 2 | passed
|  |  | UTICA-23161 - Streaming Tests for Quebec HA | EA_Quebec_01 - Utica Features on Quebec HA | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-19401 - When the WiCROS balance control is set to an equal mix of the streamed audio sig... | UTICA-23148 - Streaming Tests for Paris HA | EA_Paris_01 - Utica Features on Paris HA_EA DVT | Executed | 3 | passed
|  |  | UTICA-23161 - Streaming Tests for Quebec HA | EA_Quebec_01 - Utica Features on Quebec HA | Executed | 3 | passed
Mobile Apps | passed | MASTER-64684 - The hearing aid shall comply with MFi Accessory Interface Specification R33 Hear... | UTICA-24579 - The hearing aid shall comply with MFi Accessory Interface Specification R33 Hear... | Utica_MFi_TestRun - Utica_MFi_TestRun | Executed | N/A | passed
Regulatory | passed | MASTER-50186 - The hearing aid operations manual shall include Tinnitus Therapy feature labelin... | UTICA-24662 - Utica Feature SyRS - Regulatory | Utica - Regulatory | Executed | 1 | passed
|  |  | UTICA-24591 - Regulatory Tinnitus | Regulatory - Testing | Executed | 1 | passed