## Sheet1
Inspire Metrics Report Grand Avenue Document ID: GQ-RPT-01043 Revision: 2.0
Description of data in this document: Function: This column designates which functional team is responsible for verifying the corresponding Systems requirement. Requirement Result: This column indicates the cumulative (across final test runs) verification result for the corresponding requirement. This is the cumulative result of the corresponding Test Resullt (Column H) data. Requirement: This column lists the requirement title as shown in Polarion and provides a hyperlink to the requirement in Polarion. Test Case: This column indicates which test case(s) contain steps that verify the requirement. Please note that Polarion includes all executed tests that reference the relevant requirement. Please also note that Polarion only includes verification results in this report from the most recently executed Test Run(s). For these reasons, some results included in this report may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Run: This column lists the test run(s) that contained the test case(s) (Column D) that verify the requirement. Please note that Polarion only includes verification results in this report from the most recently executed Test Run(s), which may or may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Result Status: This column indicates the execution status for each instance of the requirement in all executed test cases. Test Step: This column indicates the location of every instance of the requirement in the corresponding Test Case (Column D). Zero-step test cases (test cases that directly pass/fail an individual requirement) will indicate N/A in this column. Test Result: This column indicates the test result for every instance of the requirement in the corresponding Test Case (Column D). This column is the input for the Requirement Result (Column B) column.
Function | Requirement Result | Requirement | Test Case | Test Run | Test Result Status | Test Step | Test Result
Systems | passed | MASTER-21791 - For the Flagship tier product, the fitting software shall provide controls for 2... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
[5 empty] | Executed | 24 | passed
Systems | passed | MASTER-21792 - For the Upper-end tier product, the fitting software shall provide controls for... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 33 | passed
[5 empty] | Executed | 50 | passed
Systems | passed | MASTER-21793 - For the Mid-end tier, the fitting software shall provide controls for 16 adjusta... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 59 | passed
[5 empty] | Executed | 76 | passed
Systems | passed | MASTER-21794 - For the Low-end tier, the fitting software shall provide controls for 12 adjusta... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 85 | passed
[5 empty] | Executed | 102 | passed
Systems | passed | MASTER-21795 - For the Entry-level tier, the fitting software shall provide controls for 10 adj... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 111 | passed
[5 empty] | Executed | 128 | passed
Systems | passed | MASTER-60779 - For hearing aids configured with a "push button" user control, the fitting softw... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 5 | passed
[5 empty] | Executed | 5 | passed
|  |  | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 16 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 32 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 17 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 129 | passed
[5 empty] | Executed | 135 | passed
[5 empty] | Executed | 137 | passed
[5 empty] | Executed | 139 | passed
[5 empty] | Executed | 141 | passed
[5 empty] | Executed | 151 | passed
[5 empty] | Executed | 157 | passed
[5 empty] | Executed | 167 | passed
[5 empty] | Executed | 168 | passed
[5 empty] | Executed | 195 | passed
[5 empty] | Executed | 212 | passed
[5 empty] | Executed | 213 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 142 | passed
[5 empty] | Executed | 148 | passed
[5 empty] | Executed | 150 | passed
[5 empty] | Executed | 152 | passed
[5 empty] | Executed | 154 | passed
[5 empty] | Executed | 164 | passed
[5 empty] | Executed | 170 | passed
[5 empty] | Executed | 203 | passed
[5 empty] | Executed | 204 | passed
[5 empty] | Executed | 210 | passed
[5 empty] | Executed | 231 | passed
[5 empty] | Executed | 248 | passed
[5 empty] | Executed | 249 | passed
Systems | passed | MASTER-60784 - In a binaural fitting session, when the Manual Power Off feature is assigned to... | MASTER-34373 - SV002E - User Control - Power Management | IP9 - General IP9 Testing & Regression | Executed | 17 | passed
[5 empty] | Executed | 20 | passed
Systems | passed | MASTER-60785 - When a binaural fitting session is initiated and the assignment of the Manual Po... | MASTER-34373 - SV002E - User Control - Power Management | IP9 - General IP9 Testing & Regression | Executed | 41 | passed
Systems | passed | MASTER-60782 - In a binaural fitting session, when the Manual Alert feature is assigned to or r... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
[5 empty] | Executed | 7 | passed
Systems | passed | MASTER-60783 - For hearing aids configured with a "rocker switch" user control, the fitting sof... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
|  |  | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 33 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 36 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 63 | passed
[5 empty] | Executed | 65 | passed
Systems | passed | MASTER-60786 - In a binaural fitting session,when the Manual Power Off feature is assigned to o... | MASTER-34373 - SV002E - User Control - Power Management | IP9 - General IP9 Testing & Regression | Executed | 17 | passed
[5 empty] | Executed | 20 | passed
Systems | passed | MASTER-60788 - When a binaural fitting session is initiated and the assignment of the Manual Po... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 34 | passed
Systems | passed | MASTER-60787 - In a binaural fitting session, when the Manual Alert feature is assigned to or r... | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
[5 empty] | Executed | 7 | passed
Systems | passed | MASTER-60791 - For hearing aids configured with an IMU, the fitting software shall have a metho... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 5 | passed
[5 empty] | Executed | 5 | passed
|  |  | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-60789 - For WiCROS Partner hearing aids configured with an IMU, the fitting software sha... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 40 | passed
|  |  | MASTER-53167 - SV246 - NFMI WiCROS - User Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 100 | passed
Systems | passed | MASTER-60792 - For hearing aids configured with a digital rotary user control, the fitting soft... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 35 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 182 | passed
Systems | passed | MASTER-60795 - In a non-WiCROS, binaural wireless fitting session, the fitting software shall h... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 7 | passed
[5 empty] | Executed | 7 | passed
|  |  | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 9 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 84 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 169 | passed
[5 empty] | Executed | 195 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 205 | passed
[5 empty] | Executed | 231 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
[5 empty] | Executed | 7 | passed
Systems | passed | MASTER-60796 - In a non-WiCROS, binaural wireless fitting session of hearing aids configured wi... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 37 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 205 | passed
Systems | passed | MASTER-60793 - In a binaural wireless fitting session, when the Manual Power Off feature is ass... | MASTER-34373 - SV002E - User Control - Power Management | IP9 - General IP9 Testing & Regression | Executed | 18 | passed
Systems | passed | MASTER-60794 - In a non-WiCROS, binaural wireless fitting session, when the Multi-Memory featur... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 36 | passed
Systems | passed | MASTER-60797 - In a non-WiCROS, binaural wireless fitting session, when the Autophone or Autoco... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 38 | passed
Systems | passed | MASTER-60798 - In a binaural wireless fitting session, when the Personal Voice Assistant featur... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 39 | passed
Systems | passed | MASTER-60926 - In a binaural wireless fitting session, when the On-Demand Adaptive Tuning featu... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 6 | passed
[5 empty] | Executed | 6 | passed
Systems | passed | MASTER-21877 - The fitting software shall have a method to individually enable and disable Memo... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 6 | passed
Systems | passed | MASTER-21878 - The fitting software shall have a method to individually assign a memory environ... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 7 | passed
Systems | passed | MASTER-21879 - The fitting software shall have a method to independently include Memory 4 and M... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 8 | passed
Systems | passed | MASTER-20651 - The manufacturing software shall enable Memories 5 and 6 in the hearing aid. | MASTER-66945 - SV121B - Manufacturing Suite (For 2.4 Devices Only) | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 38 | passed
[5 empty] | Executed | 38 | passed
Systems | passed | MASTER-20652 - The fitting software shall preclude changing the enabled/disabled state of Memor... | MASTER-42122 - SV169 - Mobile Application - Custom Memories | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 11 | passed
Systems | passed | MASTER-20657 - When the hearing aid is configured with 4 memory locations, the fitting software... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 10 | passed
Systems | passed | MASTER-20695 - The fitting software shall assign the hearing aid a Volume Control Range of 20 d... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 27 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 22 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 23 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 36 | passed
[5 empty] | Executed | 185 | passed
Systems | passed | MASTER-61649 - For standard style products, the fitting software shall have a method for assign... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 40 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 21 | passed
Systems | passed | MASTER-21884 - For custom style products, the fitting software shall have a method for assignin... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 23 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 20 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 10 | passed
[5 empty] | Executed | 14 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 183 | passed
Systems | passed | MASTER-21885 - When a Volume Control feature is enabled, the fitting software shall provide a v... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 23 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 20 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 21 | passed
Systems | passed | MASTER-21886 - When a Volume Control feature is enabled, the fitting software shall provide a v... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 26 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 22 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 23 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 36 | passed
Systems | passed | MASTER-20698 - For a Volume Step Size of 4 dB, the fitting software shall have a method for ass... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 25 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 21 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 22 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 29 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 35 | passed
Systems | passed | MASTER-64389 - For a Volume Step Size of 2 dB, the fitting software shall have a method for ass... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 42 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 24 | passed
[5 empty] | Executed | 25 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 37 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 184 | passed
Systems | passed | MASTER-21893 - When a Volume Control feature is enabled, the fitting software shall provide a v... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 34 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 29 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 30 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 39 | passed
Systems | passed | MASTER-21894 - When a Volume Control feature is enabled, the fitting software shall provide a v... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 30 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 25 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 26 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 38 | passed
Systems | passed | MASTER-21895 - When a Volume Control feature is enabled, the fitting software shall provide a v... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 29 | passed
[5 empty] | Executed | 32 | passed
|  |  | MASTER-73416 - SV002A(i)-User Control- Volume Control - Rocker Switch | IP6 - General IP6 Testing | Executed | 24 | passed
[5 empty] | Executed | 27 | passed
|  |  | MASTER-73417 - SV002A(ii)-User Control- Volume Control - Push Button | IP6 - General IP6 Testing | Executed | 25 | passed
[5 empty] | Executed | 28 | passed
|  |  | MASTER-73418 - SV002A(iii)-User Control-Volume Control-Digital Rotary | IP6 - General IP6 Testing | Executed | 37 | passed
Systems | passed | MASTER-21899 - The fitting software shall have a method to enable and disable Reset Levels on M... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 230 | passed
[5 empty] | Executed | 232 | passed
Systems | passed | MASTER-21900 - In a binaural fitting session, the fitting software shall enable or disable Rese... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 238 | passed
Systems | passed | MASTER-21901 - When a binaural pair of hearing aids is brought into a fitting session, if Reset... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 237 | passed
[5 empty] | Executed | 238 | passed
Systems | passed | MASTER-21904 - When any amount of Reserve Gain is desired, and it is not available in all chann... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 258 | passed
Systems | passed | MASTER-21909 - When Auto Acclimatization is enabled and the volume range is limited in some cha... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 163 | passed
Systems | passed | MASTER-21912 - The fitting software shall have a * Normal environment. Note: At least one envir... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 9 | passed
Systems | passed | MASTER-21914 - The fitting software shall have a Meeting environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 10 | passed
Systems | passed | MASTER-21916 - The fitting software shall have a Crowd environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 11 | passed
Systems | passed | MASTER-21918 - The fitting software shall have a Restaurant environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 12 | passed
Systems | passed | MASTER-21920 - The fitting software shall have a Car environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 13 | passed
Systems | passed | MASTER-21922 - The fitting software shall have a Auditorium environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 14 | passed
Systems | passed | MASTER-21924 - The fitting software shall have a Music environment. Note: Kiev and newer produc... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 15 | passed
Systems | passed | MASTER-21926 - The fitting software shall have a Television environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 16 | passed
Systems | passed | MASTER-21928 - The fitting software shall have a Outdoors environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 17 | passed
Systems | passed | MASTER-21930 - The fitting software shall have a Mute environment. | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 18 | passed
Systems | passed | MASTER-21931 - For the Mute environment, the fitting software shall set the gain in each MECO c... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 19 | passed
Systems | passed | MASTER-21935 - The fitting software shall have a Telephone environment. Note: The "Telephone" m... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 21 | passed
Systems | passed | MASTER-21937 - The fitting software shall have an Autophone environment. Note: The "Autophone"... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 22 | passed
Systems | passed | MASTER-21939 - For hearing aids configured with an induction coil, the fitting software shall h... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 23 | passed
Systems | passed | MASTER-21941 - For hearing aids configured with an induction coil, the fitting software shall h... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 24 | passed
Systems | passed | MASTER-21943 - The fitting software shall have an Offset (0 dB) environment. Note: The Offset e... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 25 | passed
Systems | passed | MASTER-21945 - The fitting software shall have an Offset (-6 dB) environment. Note: The Offset... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 26 | passed
Systems | passed | MASTER-21947 - The fitting software shall have an Offset (-40 dB) environment. Note: The Offset... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 27 | passed
Systems | passed | MASTER-21949 - The fitting software shall have a Receive Signal environment. Note: The Receive... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 28 | passed
Systems | passed | MASTER-21951 - For hearing aids configured with an induction coil, the fitting software shall h... | MASTER-39838 - SV206 - Subsystem Verification - SV057 - Memories in Inspire | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 29 | passed
Systems | passed | MASTER-21975 - For each music-adaptation memory, the fitting software shall have a method to en... | MASTER-34370 - SV141 - Music AEC-Inspire Interface | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
Systems | passed | MASTER-21980 - For each music-adaptation memory, the fitting software shall set the Expansion R... | MASTER-34370 - SV141 - Music AEC-Inspire Interface | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
Systems | passed | MASTER-21982 - For each memory in which Music Adaptation is enabled, the fitting software shall... | MASTER-34370 - SV141 - Music AEC-Inspire Interface | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 23 | passed
[5 empty] | Executed | 25 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
Systems | passed | MASTER-21983 - The fitting software shall determine the Music Adaptation Target Offset in dB fo... | MASTER-34370 - SV141 - Music AEC-Inspire Interface | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 25 | passed
Systems | passed | MASTER-21984 - When the professional performs a binaural fitting on two E2E Partners capable of... | MASTER-26725 - SV002A - User Control - Volume Control | IP5 - General IP5 Testing | Executed | 33 | passed
|  |  | MASTER-34370 - SV141 - Music AEC-Inspire Interface | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
Systems | passed | MASTER-55413 - When in a WiCROS fitting session, the fitting software shall enable E2E Music Ad... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 156 | passed
Systems | passed | MASTER-21987 - For each machine-adaptation memory, the fitting software shall have a method to... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
Systems | passed | MASTER-21991 - For the Flagship tier, when at least one memory location has Machine Adaptation... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 4 | passed
[5 empty] | Executed | 5 | passed
Systems | passed | MASTER-21992 - For the High-end, Mid-end, Low-end, and Entry-level tiers, when at least one mem... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 4 | passed
[5 empty] | Executed | 5 | passed
Systems | passed | MASTER-21993 - When the professional performs a binaural fitting on two E2E Partners capable of... | MASTER-28276 - SV054 - Machine Noise in Inspire | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
Systems | passed | MASTER-55414 - When in a WiCROS fitting session, the fitting software shall enable E2E Machine... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 156 | passed
Systems | passed | MASTER-21996 - For each wind-adaptation memory, the fitting software shall have a method to ena... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 11 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 38 | passed
Systems | passed | MASTER-22000 - For the Flagship tier, when at least one memory location has Wind Adaptation ena... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 10 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 10 | passed
[5 empty] | Executed | 11 | passed
Systems | passed | MASTER-22001 - For the High-end, Mid-end, Low-end, and Entry-level tiers, when at least one mem... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 10 | passed
[5 empty] | Executed | 10 | passed
Systems | passed | MASTER-22002 - When the professional performs a binaural fitting on two E2E Partners capable of... | MASTER-33351 - SV053 - Wind Adaptation | IP7 - General IP7 Testing | Executed | 11 | passed
[5 empty] | Executed | 38 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 38 | passed
Systems | passed | MASTER-55415 - When in a WiCROS fitting session, the fitting software shall enable E2E Wind Ada... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 156 | passed
Systems | passed | MASTER-50255 - The fitting software shall support the available combined Monaural Noise Reducti... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 9 | passed
Systems | passed | MASTER-22011 - The fitting software shall allow only the Static Omnidirectional Mode to be conf... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 119 | passed
Systems | passed | MASTER-22012 - The fitting software shall allow the following directionality modes to be config... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 118 | passed
Systems | passed | MASTER-22013 - When a directional mode is selected, the fitting software shall enable and disab... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 122 | passed
[5 empty] | Executed | 123 | passed
[5 empty] | Executed | 124 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 126 | passed
Systems | passed | MASTER-22017 - When Dynamic Directionality or Adaptive Directionality mode is enabled, the fitt... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 148 | passed
[5 empty] | Executed | 150 | passed
[5 empty] | Executed | 170 | passed
[5 empty] | Executed | 172 | passed
[5 empty] | Executed | 192 | passed
[5 empty] | Executed | 194 | passed
Systems | passed | MASTER-22018 - When Dynamic Directionality w/ HF Directionality or Adaptive Directionality w/ H... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 160 | passed
[5 empty] | Executed | 162 | passed
[5 empty] | Executed | 182 | passed
[5 empty] | Executed | 184 | passed
[5 empty] | Executed | 204 | passed
[5 empty] | Executed | 206 | passed
Systems | passed | MASTER-22023 - The fitting software shall have a method to collectively enable and disable E2E... | MASTER-35564 - AV198: High Frequency Directionality/Immersion Directionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
Systems | passed | MASTER-55416 - When in a WiCROS fitting session, the fitting software shall enable E2E Directio... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 156 | passed
Systems | passed | MASTER-49937 - In a binaural wireless fitting session, the fitting software shall ensure that t... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 113 | passed
Systems | passed | MASTER-49934 - When a binaural wireless fitting session is entered, and the Directionality Mode... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 120 | passed
Systems | passed | MASTER-49935 - Upon selection of which hearing aid has the correct Directionality Mode and the... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 123 | passed
Systems | passed | MASTER-22025 - When the Directionality Mode is set to any of the following, the fitting softwar... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 78 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-51789 - The fitting software shall set the Expansion Ratio in each MECO channel in the h... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 21 | passed
Systems | passed | MASTER-22029 - For a hearing aid that has the Autophone or Autocoil memory environment enabled,... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 22 | passed
[5 empty] | Executed | 98 | passed
Systems | passed | MASTER-22030 - The fitting software shall have a method to enable and disable Microphone/Teleco... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 29 | passed
[5 empty] | Executed | 59 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 105 | passed
Systems | passed | MASTER-22031 - The fitting software shall have a method to select and program the M/T Microphon... | MASTER-34301 - SV014 - Telephone Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 60 | passed
[5 empty] | Executed | 73 | passed
[5 empty] | Executed | 118 | passed
Systems | passed | MASTER-60800 - The fitting software shall automatically enable, automatically disable, and prov... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 10 | passed
[5 empty] | Executed | 10 | passed
|  |  | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 24 | passed
[5 empty] | Executed | 67 | passed
[5 empty] | Executed | 68 | passed
[5 empty] | Executed | 80 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 13 | passed
[5 empty] | Executed | 15 | passed
[5 empty] | Executed | 20 | passed
[5 empty] | Executed | 22 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 34 | passed
[5 empty] | Executed | 37 | passed
[5 empty] | Executed | 39 | passed
[5 empty] | Executed | 64 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 67 | passed
[5 empty] | Executed | 71 | passed
[5 empty] | Executed | 80 | passed
[5 empty] | Executed | 85 | passed
[5 empty] | Executed | 88 | passed
[5 empty] | Executed | 100 | passed
[5 empty] | Executed | 105 | passed
[5 empty] | Executed | 111 | passed
[5 empty] | Executed | 123 | passed
[5 empty] | Executed | 153 | passed
[5 empty] | Executed | 161 | passed
[5 empty] | Executed | 163 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 19 | passed
[5 empty] | Executed | 20 | passed
[5 empty] | Executed | 46 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 8 | passed
[5 empty] | Executed | 8 | passed
|  |  | MASTER-26428 - SV061B - Low Battery - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 5 | passed
[5 empty] | Executed | 7 | passed
Systems | passed | MASTER-22090 - In a binaural fitting, the fitting software shall provide a method to manually e... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 202 | passed
[5 empty] | Executed | 203 | passed
Systems | passed | MASTER-60810 - The fitting software shall automatically configure the audible indicator stimulu... | MASTER-47282 - SV002F - User Controls - Start/Stop Stream | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 67 | passed
|  |  | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 26 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 58 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 81 | passed
[5 empty] | Executed | 89 | passed
[5 empty] | Executed | 112 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 12 | passed
[5 empty] | Executed | 21 | passed
[5 empty] | Executed | 47 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 8 | passed
[5 empty] | Executed | 8 | passed
Systems | passed | UTICA-24565 - When a memory environment is assigned, the fitting software shall assign the def... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 38 | passed
Systems | passed | UTICA-24845 - For any enabled memory not assigned with a telephone memory environment, the fit... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 39 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-60809 - For any enabled memory assigned with a telephone memory environment, the fitting... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 13 | passed
Systems | passed | MASTER-60812 - When an indicator is configured as a tone indicator, the fitting software shall... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 27 | passed
[5 empty] | Executed | 42 | passed
[5 empty] | Executed | 70 | passed
[5 empty] | Executed | 90 | passed
Systems | passed | MASTER-22115 - When both the Volume and Tinnitus indicator groups are enabled, the fitting soft... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 36 | passed
Systems | passed | MASTER-60816 - When both the Volume and Accessory Volume indicator groups are enabled, the fitt... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 165 | passed
Systems | passed | MASTER-21034 - When a speech indicator is assigned to a memory, the fitting software shall set... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
Systems | passed | MASTER-21035 - When a tone indicator is assigned to a memory, the fitting software shall set th... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 69 | passed
Systems | passed | MASTER-60815 - In a binaural wireless fitting session, the fitting software shall apply indicat... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 147 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 30 | passed
Systems | passed | MASTER-60818 - When a binaural wireless fitting session is entered and the indicator stimulus c... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 197 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 27 | passed
Systems | passed | MASTER-60819 - Upon selection of which hearing aid has the correct indicator stimulus configura... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 198 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 28 | passed
Systems | passed | MASTER-60822 - When an indicator group is configured as speech indicator(s), the fitting softwa... |  |  | Executed | 49 | passed
Systems | passed | MASTER-60827 - In a binaural wireless fitting, the fitting software shall apply Cue indicator s... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 61 | passed
Systems | passed | MASTER-60828 - When a binaural wireless programming session is entered and the Cue indicator se... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 59 | passed
Systems | passed | MASTER-60825 - Upon selection of which hearing aid has the correct Cue indicator setting, the f... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 60 | passed
Systems | passed | MASTER-60826 - The fitting software shall have a method to independently set the base presentat... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 141 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 75 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
Systems | passed | MASTER-60829 - When individual indicators within an indicator group can be configured independe... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 142 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 14 | passed
Systems | passed | MASTER-61648 - The fitting software shall have a method to set the base presentation level for... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 73 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 13 | passed
Systems | passed | MASTER-61647 - The fitting software shall have a method to set the base presentation level for... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 13 | passed
Systems | passed | MASTER-60831 - In a binaural fitting, the fitting software shall provide a method to adjust the... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 143 | passed
[5 empty] | Executed | 144 | passed
|  |  | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 74 | passed
|  |  | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 13 | passed
[5 empty] | Executed | 14 | passed
[5 empty] | Executed | 15 | passed
[5 empty] | Executed | 16 | passed
Systems | passed | MASTER-60830 - When the base presentation level for a speech indicator has been adjusted from t... | MASTER-28187 - SV186C - Indicator Level | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 92 | passed
Systems | passed | MASTER-22138 - The fitting software shall have a method to enable and disable level dependence... | MASTER-39757 - AV186C - Indicator Level | IP11 - Bug Fixes & Regression | Executed | 16 | passed
[5 empty] | Executed | 25 | passed
Systems | passed | MASTER-60835 - The fitting software shall have a method to select the language and gender of th... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 32 | passed
Systems | passed | MASTER-60833 - When language or gender of the speech indicator is changed, the fitting software... | MASTER-28186 - SV186B - Improved Indicators - Voice | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 31 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-22145 - The fitting software shall have a method to demonstrate each indicator on the he... | MASTER-28185 - SV186A - Improved Tone Indicators | IP7 - General IP7 Testing | Executed | 195 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
[5 empty] | Executed | 9 | passed
Systems | passed | MASTER-55718 - In a WiCROS fitting session, the fitting software shall have a method to indepen... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 54 | passed
Systems | passed | MASTER-22147 - The fitting software shall have a method to enable and disable DTMF Remote Contr... | MASTER-75313 - SV097 - DTMF | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 15 | passed
[5 empty] | Executed | 208 | passed
Systems | passed | MASTER-22148 - If a memory is configured with the microphone off and the DTMF Remote Control en... | MASTER-75313 - SV097 - DTMF | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 250 | passed
Systems | passed | MASTER-22151 - The fitting software shall have a method to enable and disable DTMF Remote Adjus... | MASTER-75313 - SV097 - DTMF | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 15 | passed
Systems | passed | MASTER-22152 - The fitting software shall have a method to change the last two digits of the DT... | MASTER-75313 - SV097 - DTMF | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 74 | passed
[5 empty] | Executed | 156 | passed
Systems | passed | MASTER-22154 - The fitting software shall provide an interface for assigning a short, medium, o... | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 8 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 30 | passed
Systems | passed | MASTER-22155 - The fitting software shall set the power on delay in the hearing aid according t... | MASTER-26717 - SV070 - Power On Delay | IP6 - General IP6 Testing | Executed | 11 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 30 | passed
Systems | passed | MASTER-22160 - The fitting software shall display the discovered wireless hearing aids within r... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 7 | passed
[5 empty] | Executed | 8 | passed
Systems | passed | MASTER-22161 - The fitting software shall display the battery level for each discovered hearing... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 7 | passed
Systems | passed | MASTER-22162 - The fitting software shall have a method for selecting a hearing aid for use in... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 7 | passed
Systems | passed | MASTER-22163 - The fitting software shall have a method to designate a hearing aid as left or r... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 8 | passed
Systems | passed | MASTER-22164 - The fitting software shall have a method to designate two hearing aids as a bina... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 8 | passed
Systems | passed | MASTER-21159 - The fitting software shall have a method to assign one Favorite Button function... | MASTER-42279 - SV187 - 2.4GHz -Remote Control - General Functionality | Continuation 1 - General Testing for Utica Continuation IP1 | Executed | 211 | passed
[5 empty] | Executed | 216 | passed
[5 empty] | Executed | 225 | passed
[5 empty] | Executed | 227 | passed
[5 empty] | Executed | 236 | passed
[5 empty] | Executed | 245 | passed
[5 empty] | Executed | 255 | passed
[5 empty] | Executed | 257 | passed
Systems | passed | MASTER-22174 - When two hearing aids are E2E Partners, the fitting software shall assign the sa... | MASTER-34404 - SV105B - Noise Control - 2.4GHz | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 12 | passed
Systems | passed | UTICA-24846 - The fitting software shall have a method to set the Initial Audio Streaming for... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Amplifon Testing - Utica Continuation | Executed | 17 | passed
Systems | passed | MASTER-42077 - In a binaural fitting on two E2E Partners, the fitting software shall have a met... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 8 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 167 | passed
[5 empty] | Executed | 169 | passed
Systems | passed | MASTER-22196 - The fitting software shall preclude enabling of Initiate E2E Phone Streaming in... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 8 | passed
Systems | passed | MASTER-22197 - The fitting software shall not allow both E2E Partners to have Initiate E2E Phon... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 40 | passed
Systems | passed | MASTER-22198 - The fitting software shall allow Initiate E2E Phone Streaming to be enabled in o... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 40 | passed
Systems | passed | MASTER-22200 - When Initiate E2E Phone Streaming is enabled on one hearing aid, the fitting sof... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 15 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 43 | passed
Systems | passed | MASTER-22201 - When a memory location that has been configured with the Receive E2E Phone Strea... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 18 | passed
[5 empty] | Executed | 43 | passed
Systems | passed | MASTER-22202 - When Initiate E2E Phone Streaming is disabled on one hearing aid, the fitting so... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 169 | passed
Systems | passed | MASTER-22204 - When a binaural programming session is entered and the E2E Phone Streaming setti... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 176 | passed
[5 empty] | Executed | 178 | passed
Systems | passed | MASTER-22205 - Upon selection of which hearing aid has the correct E2E Phone Streaming settings... | MASTER-26384 - SV010 - E2E Phone Streaming | Z22A - New Rechargeable Battery | Executed | 178 | passed
Systems | passed | MASTER-22208 - The fitting software shall be able to detect and program a WiCROS Transmitter. | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 6 | passed
Systems | passed | MASTER-22209 - The fitting software shall allow a WiCROS Transmitter to be programmed only as a... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 7 | passed
Systems | passed | MASTER-22210 - When the fitting software detects a WiCROS Transmitter and a WiCROS-capable hear... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 9 | passed
Systems | passed | MASTER-22211 - The fitting software shall only allow a WiCROS Transmitter to be programmed as p... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 7 | passed
Systems | passed | MASTER-22212 - The fitting software shall have a method to remove the WiCROS Partner configurat... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 14 | passed
Systems | passed | MASTER-22214 - The fitting software shall allow programming of a WiCROS Partner hearing aid whe... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 16 | passed
Systems | passed | MASTER-22216 - The fitting software shall allow entry into a WiCROS fitting session only when t... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 20 | passed
Systems | passed | MASTER-22219 - For each memory location in the WiCROS Partner hearing aid which has been config... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 27 | passed
Systems | passed | MASTER-22220 - For each memory location in the WiCROS Partner hearing aid which has been config... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 26 | passed
Systems | passed | MASTER-22223 - When the BiCROS configuration is selected, the fitting software shall have a bal... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 31 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-22227 - When the CROS or BiCROS configuration has been selected, the fitting software sh... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 11 | passed
[5 empty] | Executed | 39 | passed
Systems | passed | MASTER-22228 - When the CROS or BiCROS configuration has been selected, the fitting software sh... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 39 | passed
Systems | passed | MASTER-22238 - When both WiCROS Partners or only a WiCROS Partner hearing aid is connected, the... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 90 | passed
[5 empty] | Executed | 92 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 94 | passed
[5 empty] | Executed | 96 | passed
Systems | passed | MASTER-22251 - The fitting software shall preclude enabling the E2E Phone Streaming feature on... | MASTER-53147 - SV244 - NFMI WiCROS - Inspire Session | IP10 - Bug Fixes & Regression | Executed | 43 | passed
Systems | passed | MASTER-22253 - The fitting software shall not display the logged data from the WiCROS Transmitt... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 136 | passed
Systems | passed | MASTER-50682 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 3 minutes afte... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Amplifon Testing - Utica Continuation | Executed | 15 | passed
Systems | passed | MASTER-50677 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 3 minutes afte... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Amplifon Testing - Utica Continuation | Executed | 8 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 14 | passed
Systems | passed | MASTER-50674 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 3 minutes afte... | MASTER-47234 - SV125B - 2.4ghz Audio Streaming and accessories | Amplifon Testing - Utica Continuation | Executed | 8 | passed
[5 empty] | Executed | 11 | passed
[5 empty] | Executed | 14 | passed
Systems | passed | MASTER-50675 - When the fitting software sends a pairing command to the connected hearing aids,... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 93 | passed
[5 empty] | Executed | 115 | passed
Systems | passed | MASTER-50680 - If no additional accessories have been paired after 30 seconds, the fitting soft... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 88 | passed
[5 empty] | Executed | 110 | passed
Systems | passed | MASTER-50681 - Upon receipt of a pairing request from a 2.4 GHz MFA greater than 3 minutes afte... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 101 | passed
[5 empty] | Executed | 123 | passed
Systems | passed | MASTER-50678 - In a binaural fitting session, the fitting software shall send a pairing deletio... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 123 | passed
Systems | passed | MASTER-50679 - When the fitting software sends a pairing deletion command to the connected hear... | MASTER-49134 - SV233 - MFA Pairing | Model S IP 6-5 | Executed | 101 | passed
[5 empty] | Executed | 123 | passed
Systems | passed | MASTER-22269 - The fitting software shall have a method to initiate a wireless fitting session... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 7 | passed
[5 empty] | Executed | 8 | passed
Systems | passed | MASTER-22270 - The fitting software shall have a method to initiate a wired fitting session wit... | MASTER-34423 - SV153 - Wired HA Functionality | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 7 | passed
Systems | passed | MASTER-21278 - When the professional starts a fitting session, the fitting software shall set t... | MASTER-34256 - SV008 - Inspire Session: Features Disabled, End Session, Session Timeout | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 35 | passed
[5 empty] | Executed | 84 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 168 | passed
[5 empty] | Executed | 213 | passed
Systems | passed | MASTER-22272 - The fitting software shall have a method to end a fitting session. | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 13 | passed
Systems | passed | MASTER-22273 - When a wireless fitting session is ended, the fitting software shall send a comm... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 14 | passed
Systems | passed | MASTER-22275 - The fitting software shall have a method to perform a field firmware update. | MASTER-34468 - SV069 - Verify FFU | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 46 | passed
Systems | passed | MASTER-22276 - The fitting software shall allow field firmware updates to be targeted to specif... | MASTER-34468 - SV069 - Verify FFU | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 13 | passed
[5 empty] | Executed | 46 | passed
Systems | passed | MASTER-55719 - The fitting software shall support programming Demo hearing aids. | MASTER-56940 - SV251 - Demo Hearing aid - Inspire | Continuation 2 - General Testing for Utica Continuation 2 | Executed | 27 | passed
Systems | passed | MASTER-55727 - When updating the product tier in a Demo hearing aid, the fitting software shall... | MASTER-56940 - SV251 - Demo Hearing aid - Inspire | Continuation 2 - General Testing for Utica Continuation 2 | Executed | 32 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-55728 - When updating the product tier in a Demo hearing aid, the fitting software shall... | MASTER-56940 - SV251 - Demo Hearing aid - Inspire | Continuation 2 - General Testing for Utica Continuation 2 | Executed | 32 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-22278 - The fitting software shall utilize the patient's audiometric configuration to se... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 2 | passed
|  |  | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
[5 empty] | Executed | 12 | passed
[5 empty] | Executed | 15 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 26 | passed
Systems | passed | MASTER-22279 - For each memory location in the hearing aid configured with a memory environment... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 3 | passed
Systems | passed | MASTER-41122 - The fitting software shall preclude enabling Frequency Translation in the Music... | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 9 | passed
Systems | passed | MASTER-22280 - The fitting software shall provide a visual indication when Frequency Translatio... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 4 | passed
Systems | passed | MASTER-22281 - For each memory location in the hearing aid, the fitting software shall have a m... | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 17 | passed
[5 empty] | Executed | 18 | passed
[5 empty] | Executed | 19 | passed
[5 empty] | Executed | 21 | passed
Systems | passed | MASTER-22286 - The fitting software shall have a method to adjust the Splitting Frequency. | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 18 | passed
Systems | passed | MASTER-22287 - The fitting software shall have a method to adjust the Knee Frequency. | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 17 | passed
Systems | passed | MASTER-22288 - The fitting software shall have a method to adjust the Warping Ratio. | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 19 | passed
Systems | passed | MASTER-22289 - The fitting software shall have a method to adjust the Frequency Translation Gai... | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 21 | passed
Systems | passed | MASTER-22290 - The fitting software settings for Frequency Translation shall be independent of... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 5 | passed
|  |  | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 61 | passed
Systems | passed | MASTER-22291 - The fitting software shall have a method to adjust the settings for Frequency Tr... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 6 | passed
Systems | passed | MASTER-22292 - The fitting software shall provide the option to display the effects of Frequenc... | MASTER-47197 - SV229 - Subsystem Verification - SV095 - Frequency Translation | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 7 | passed
Systems | passed | MASTER-22293 - When the fitting software is in Test Mode and Full On Gain or Reference Test Gai... | MASTER-34623 - AV095 - Frequency Translation | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
[5 empty] | Executed | 34 | passed
Systems | passed | MASTER-22294 - When the fitting software is in Test Mode and User Gain is selected, the fitting... | MASTER-34622 - SV095 - Frequency Translation | IP9 - General IP9 Testing & Regression | Executed | 66 | passed
Systems | passed | MASTER-61650 - For standard style products, the fitting software shall have a method to adjust... | UTICA-9340 - Utica User Controls (Test By Analysis) | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 41 | passed
Systems | passed | MASTER-22345 - The fitting software shall have a method to enable and disable the microphone du... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 369 | passed
[5 empty] | Executed | 371 | passed
[5 empty] | Executed | 377 | passed
Systems | passed | MASTER-22352 - For each memory location in the Flagship and Upper-end tier hearing aid that has... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 53 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 79 | passed
|  |  | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-22353 - For each memory location in the Low-end and Entry-level tier hearing aid that ha... | MASTER-34406 - AV119 - Bands and Channels | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 87 | passed
[5 empty] | Executed | 105 | passed
[5 empty] | Executed | 113 | passed
[5 empty] | Executed | 131 | passed
|  |  | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-22354 - For each memory location in the Flagship and Upper-end tier hearing aid that has... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-22355 - For each memory location in the Low-end and Entry-level tier hearing aid that ha... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
[5 empty] | Executed | 69 | passed
[5 empty] | Executed | 82 | passed
[5 empty] | Executed | 93 | passed
Systems | passed | MASTER-22357 - The fitting software Stimulus settings shall be such that the maximum 22 kHz-ban... | MASTER-34428 - AV140 - Tinnitus - Acoustic Measurements | Utica Automation - Automation for 2.4 GHz Product Verification | Executed | 99 | passed
[5 empty] | Executed | 100 | passed
[5 empty] | Executed | 99 | passed
[5 empty] | Executed | 100 | passed
Systems | passed | MASTER-21398 - For memory locations where the fitting software is allowed to enable Tinnitus Th... | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 285 | passed
[5 empty] | Executed | 286 | passed
[5 empty] | Executed | 288 | passed
[5 empty] | Executed | 289 | passed
Systems | passed | MASTER-64689 - The fitting software shall provide instructions on how to utilize the Tinnitus S... | MASTER-64571 - SV140B - Tinnitus Stimulus Personalization (Wizard) | IP7 - General IP7 Testing | Executed | 59 | passed
Systems | passed | MASTER-64687 - When an attempt is made to exit Tinnitus Stimulus Customization, the fitting sof... | MASTER-64571 - SV140B - Tinnitus Stimulus Personalization (Wizard) | IP7 - General IP7 Testing | Executed | 46 | passed
Systems | passed | MASTER-64694 - Upon completing the MDL Test for all three frequency groups, the fitting softwar... | MASTER-64571 - SV140B - Tinnitus Stimulus Personalization (Wizard) | IP7 - General IP7 Testing | Executed | 28 | passed
Systems | passed | MASTER-64697 - Upon completing the MML test, the fitting software shall utilize the MDL and MML... | MASTER-64571 - SV140B - Tinnitus Stimulus Personalization (Wizard) | IP7 - General IP7 Testing | Executed | 10 | passed
[5 empty] | Executed | 49 | passed
[5 empty] | Executed | 60 | passed
[5 empty] | Executed | 61 | passed
Systems | passed | MASTER-22368 - The fitting software shall have a method to enable and disable Auto Acclimatizat... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 15 | passed
Systems | passed | MASTER-22369 - The fitting software shall have a method to select one of the following starting... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 17 | passed
Systems | passed | MASTER-22370 - The fitting software shall set the ending setting to Level 3 (final gain setting... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 72 | passed
Systems | passed | MASTER-22372 - The fitting software shall have a method to select one of the following duration... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 16 | passed
Systems | passed | MASTER-64400 - The fitting software shall have a method to select one of the following duration... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 16 | passed
Systems | passed | MASTER-22373 - The fitting software shall allow adjustments to be made to the gain settings at... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 21 | passed
Systems | passed | MASTER-22374 - When adjustments to the gain settings are made after the Auto Acclimatization st... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 79 | passed
Systems | passed | MASTER-22375 - The fitting software shall have a method to program the hearing aid to the Auto... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 13 | passed
Systems | passed | MASTER-22377 - When Level 1 starting settings have been selected, the fitting software shall se... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 37 | passed
Systems | passed | MASTER-22378 - When Level 2 starting settings have been selected, the fitting software shall se... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 37 | passed
Systems | passed | MASTER-22380 - For each frequency band in the hearing aid, the fitting software shall determine... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 66 | passed
Systems | passed | MASTER-22381 - The fitting software shall determine the Auto Acclimatization Time Per Step in t... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 76 | passed
Systems | passed | MASTER-22382 - The fitting software shall set the Auto Acclimatization Time Per Step in the hea... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 16 | passed
Systems | passed | MASTER-22384 - When Auto Acclimatization is in progress, the fitting software shall display how... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 76 | passed
Systems | passed | MASTER-22385 - When Auto Acclimatization is in progress, the fitting software shall display the... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 75 | passed
Systems | passed | MASTER-22386 - When Auto Acclimatization is in progress, the fitting software shall display the... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 75 | passed
Systems | passed | MASTER-22387 - The fitting software shall allow adjustments to be made to the gain settings whi... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 79 | passed
Systems | passed | MASTER-22388 - When adjustments to the gain settings are made while Auto Acclimatization adapta... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 79 | passed
Systems | passed | MASTER-22389 - The fitting software shall have a method to disable the Auto Acclimatization ada... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 89 | passed
Systems | passed | MASTER-22390 - When Auto Acclimatization is disabled prior to completion of adaptation, the fit... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 91 | passed
Systems | passed | MASTER-22391 - When Auto Acclimatization is disabled prior to completion of adaptation, the fit... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 90 | passed
Systems | passed | MASTER-22392 - When Auto Acclimatization has been disabled prior to completion of adaptation an... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 90 | passed
Systems | passed | MASTER-22393 - When Auto Acclimatization is in progress, the fitting software shall have a meth... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 80 | passed
Systems | passed | MASTER-22394 - When Auto Acclimatization has been paused, the fitting software shall have a met... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 87 | passed
Systems | passed | MASTER-22395 - When Auto Acclimatization has been paused, the fitting software shall persist th... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 86 | passed
Systems | passed | MASTER-22396 - When Auto Acclimatization is in progress, the fitting software shall have a meth... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 112 | passed
Systems | passed | MASTER-22397 - When Auto Acclimatization is reset, the fitting software shall return all Auto A... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 91 | passed
Systems | passed | MASTER-22398 - In a binaural fitting when Auto Acclimatization is enabled, the duration setting... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 104 | passed
Systems | passed | MASTER-22399 - When the Synchronize Right with Left option is selected, the fitting software sh... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 106 | passed
[5 empty] | Executed | 108 | passed
Systems | passed | MASTER-22400 - When the Synchronize Left with Right option is selected, the fitting software sh... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 119 | passed
Systems | passed | MASTER-22402 - The fitting software shall provide a visual indication when Auto Acclimatization... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 14 | passed
Systems | passed | MASTER-22403 - The fitting software shall provide undo and redo capabilities within a fitting s... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 18 | passed
Systems | passed | MASTER-22404 - When Auto Acclimatization is disabled and the feature has not been initiated in... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 20 | passed
Systems | passed | MASTER-22405 - When Auto Acclimatization is re-enabled after being disabled with the feature no... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 103 | passed
Systems | passed | MASTER-22408 - When a session connect to a saved fitting session is performed, the fitting soft... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 128 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-22411 - The fitting software shall prevent changes in gain settings for Auto Acclimatiza... | MASTER-34485 - SV072 - Auto Acclimatization | Model S IP8 | Executed | 24 | passed
Systems | passed | MASTER-22465 - The fitting software shall have a method to select one of the following experien... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 17 | passed
Systems | passed | MASTER-22466 - The fitting software shall allow the selection of any experience level, regardle... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 15 | passed
Systems | passed | MASTER-22467 - The fitting software shall allow independent experience level selections for eac... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 43 | passed
Systems | passed | MASTER-22468 - When Level 1 is selected, the fitting software shall set the hearing aid to 75%... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 70 | passed
Systems | passed | MASTER-22469 - When Level 2 is selected, the fitting software shall set the hearing aid to 85%... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 71 | passed
Systems | passed | MASTER-22470 - When Level 3 is selected, the fitting software shall set the hearing aid to 100%... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 72 | passed
Systems | passed | MASTER-22471 - When Level 4 is selected, the fitting software shall set the hearing aid to 110%... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 73 | passed
Systems | passed | MASTER-22472 - When Level T is selected, the fitting software shall set the hearing aid to the... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 44 | passed
Systems | passed | MASTER-22476 - When the hearing aid (monaural fitting) or either hearing aid (binaural fitting)... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 12 | passed
[5 empty] | Executed | 41 | passed
Systems | passed | MASTER-22477 - When the Minimal to No Experience option is selected, the fitting software shall... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 14 | passed
Systems | passed | MASTER-22478 - When the Starkey option is selected, the fitting software shall set the initial... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 48 | passed
Systems | passed | MASTER-22479 - When the Other Manufacturer option is selected, the fitting software shall have... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 50 | passed
Systems | passed | MASTER-22480 - When the Other Manufacturer option is selected and a manufacturer other than Wid... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 56 | passed
Systems | passed | MASTER-22481 - When the Other Manufacturer option is selected and a manufacturer other than Wid... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 59 | passed
[5 empty] | Executed | 62 | passed
Systems | passed | MASTER-22482 - When the Other Manufacturer option is selected and the manufacturer Widex is cho... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 51 | passed
Systems | passed | MASTER-22483 - When the Prefers Power Fit option is selected, the fitting software shall set th... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 53 | passed
Systems | passed | MASTER-22485 - The fitting software shall provide Experience Manager feature capabilities when... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 12 | passed
Systems | passed | MASTER-22486 - When the patient's birth date indicates an age of less than 18 years, the fittin... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 23 | passed
Systems | passed | MASTER-22487 - When Level 1, 2 or 4 is selected, the fitting software shall provide a visual in... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 17 | passed
Systems | passed | MASTER-22488 - The fitting software shall display the frequency response for the selected exper... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 15 | passed
Systems | passed | MASTER-22489 - The fitting software shall display the Level 3 target frequency response regardl... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 15 | passed
Systems | passed | MASTER-22490 - The fitting software shall allow adjustments to be made to the gain settings whe... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 16 | passed
Systems | passed | MASTER-22491 - When adjustments to the gain settings are made when a given experience level is... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 30 | passed
Systems | passed | MASTER-22492 - The fitting software shall apply the experience level settings to all enabled me... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 34 | passed
Systems | passed | MASTER-22493 - The fitting software shall have a method apply Best Fit to the hearing aid setti... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 36 | passed
Systems | passed | MASTER-22494 - The fitting software shall provide undo and redo capabilities within a fitting s... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 37 | passed
Systems | passed | MASTER-22502 - For a User Preference, the fitting software shall have a method to enable and di... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 76 | passed
Systems | passed | MASTER-22503 - For a User Preference, the fitting software shall have a method to select one of... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 77 | passed
Systems | passed | MASTER-22504 - For a User Preference, the fitting software shall have a method to enable and di... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 78 | passed
Systems | passed | MASTER-22506 - When a Best Fit or Target Match is performed, and the User Preference to apply t... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 84 | passed
Systems | passed | MASTER-22507 - When a Best Fit or Target Match is performed, and the User Preference to apply t... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 81 | passed
Systems | passed | MASTER-22509 - The Experience Manager feature shall be compatible with all fitting formulas use... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 88 | passed
Systems | passed | MASTER-22510 - When the fitting formula selection is changed, and the User Preference to perfor... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 89 | passed
Systems | passed | MASTER-22511 - When the fitting formula selection is changed, and the User Preference to perfor... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 93 | passed
Systems | passed | MASTER-22513 - The fitting software shall preclude Experience Manager and Auto Acclimatization... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 109 | passed
Systems | passed | MASTER-22514 - In a binaural fitting, the fitting software shall allow Experience Manager and A... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 111 | passed
Systems | passed | MASTER-22521 - When a session connect to a Milan/Greater saved fitting session is performed, th... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 103 | passed
Systems | passed | MASTER-22523 - When a fitting session is started using session connect, the fitting software sh... | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 102 | passed
Systems | passed | MASTER-22525 - The manufacturing software shall set the default experience level to Level 3. | MASTER-28184 - SV181 - Experience Manager | IP6 - General IP6 Testing | Executed | 119 | passed
Systems | passed | MASTER-51859 - The fitting software shall provide a method to enter fall risk screening results... | MASTER-73156 - SV240A Fall Management for Lower Tiers | IP7 - General IP7 Testing | Executed | 47 | passed
|  |  | MASTER-51908 - SV240 - Fall management | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 80 | passed
[5 empty] | Executed | 80 | passed
Systems | passed | MASTER-56575 - The fitting software shall display the latest results of Initiated Self Check, i... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 34 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 5 | passed
Systems | passed | MASTER-56578 - In a binaural fitting, the fitting software shall have a method to select left,... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 17 | passed
[5 empty] | Executed | 19 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 12 | passed
[5 empty] | Executed | 14 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 18 | passed
Systems | passed | MASTER-56583 - The fitting software shall have a method to command the hearing aid to perform I... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 15 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 10 | passed
Systems | passed | MASTER-56585 - Upon receipt of Initiated Self Check results from the hearing aid, the fitting s... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 43 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 22 | passed
[5 empty] | Executed | 24 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 30 | passed
[5 empty] | Executed | 32 | passed
Systems | passed | MASTER-56586 - When a session connect to a saved fitting session is performed, the fitting soft... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 76 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 35 | passed
Systems | passed | MASTER-56589 - When a factory reset is performed, the fitting software shall clear the Initiate... | MASTER-61017 - SV269 - SelfCheck 3.0, for Sydney and forward | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 66 | passed
|  |  | MASTER-56873 - SV249 - Self Check 2.0 (Inspire and Mobile) | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 37 | passed
Systems | passed | MASTER-60897 - The fitting software shall provide a method to view the data log information fro... | MASTER-34494 - SV129 - Data Log | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 25 | passed
Systems | passed | MASTER-60901 - The fitting software shall provide a method to reset the data log information in... | MASTER-61075 - SV270 - On-Demand Adaptive Tuning | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 86 | passed
[5 empty] | Executed | 87 | passed
[5 empty] | Executed | 86 | passed
[5 empty] | Executed | 87 | passed
Systems | passed | MASTER-63141 - The AutoREM feature shall support unaided measurements at an input SPL of 65 dB. | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 45 | passed
Systems | passed | MASTER-63145 - AutoREM testing shall only be performed using normal memory. | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 26 | passed
Systems | passed | MASTER-63147 - The AutoREM feature shall support all NP5 and newer hearing aids. | MASTER-64353 - SV274 - Auto REM compatibility test | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 9 | passed
Systems | passed | MASTER-63146 - The AutoREM feature shall support the following third-party REM hardware and sof... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 5 | passed
Systems | passed | MASTER-63151 - The AutoREM feature shall guide the user through the probe calibration procedure... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 28 | passed
Systems | passed | MASTER-63148 - The AutoREM feature shall explicitly show when the calibration is successful or... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 36 | passed
Systems | passed | MASTER-63149 - The AutoREM feature shall clearly indicate if the probe is outside of its allowe... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 30 | passed
[5 empty] | Executed | 32 | passed
Systems | passed | MASTER-63152 - The AutoREM feature shall inform the user if no supported measurement tool is co... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 21 | passed
Systems | passed | MASTER-63157 - The AutoREM feature shall provide on-screen user instructions at each step of th... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 38 | passed
[5 empty] | Executed | 50 | passed
Systems | passed | MASTER-63158 - The AutoREM feature shall reset the experience level to its value prior to runni... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 58 | passed
Systems | passed | MASTER-64306 - The fitting software shall indicate when the absolute difference between the ave... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 84 | passed
Systems | passed | MASTER-63155 - The AutoREM feature shall perform reasonableness checks on measurement values. | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 81 | passed
Systems | passed | MASTER-63156 - AutoREM targets shall be determined by the user-selected fitting formula. | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 59 | passed
[5 empty] | Executed | 102 | passed
Systems | passed | MASTER-63161 - The fitting software shall automatically measure and apply offsets to match the... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 51 | passed
Systems | passed | MASTER-63164 - The AutoREM feature shall store the model error in the Noah file and in the hear... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 62 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-63172 - A user shall be given the option of canceling out of the measurement process at... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 86 | passed
[5 empty] | Executed | 93 | passed
[5 empty] | Executed | 98 | passed
Systems | passed | MASTER-63175 - The AutoREM feature shall provide the user with a list of maximum sound levels r... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 52 | passed
Systems | passed | MASTER-63176 - If a sound level above the configured maximum allowed level is detected, the Aut... | MASTER-64303 - SV273 - Inspire AutoREM Test Case | Requirements Cleanup 5 - Polarion Cleanup Round 5 | Executed | 71 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-22529 - The fitting software shall have a method to initiate the SwapFit feature. | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 2 | passed
Systems | passed | MASTER-22530 - When the original hearing aid and the destination hearing aid have a feature/cap... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 3 | passed
Systems | passed | MASTER-22531 - When the destination hearing aid has a feature/capability that is not available... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 4 | passed
Systems | passed | MASTER-22532 - When the original hearing aid has a feature/capability that is not available in... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 5 | passed
Systems | passed | MASTER-22533 - In a binaural fitting session, the fitting software shall transfer data from the... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 6 | passed
Systems | passed | MASTER-22534 - The fitting software shall mute the destination hearing aid during the transfer... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 7 | passed
Systems | passed | MASTER-22535 - The fitting software shall display progress status as the SwapFit feature procee... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 8 | passed
Systems | passed | MASTER-22537 - When the destination hearing aid is a RIC, the fitting software shall provide a... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 9 | passed
Systems | passed | MASTER-22538 - The fitting software shall use the selected receiver matrix to program the desti... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 10 | passed
Systems | passed | MASTER-22540 - When the original hearing aid and the destination hearing aid have exactly the s... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 11 | passed
Systems | passed | MASTER-22541 - When the original hearing aid and the destination hearing aid have different typ... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 12 | passed
Systems | passed | MASTER-22542 - When the original hearing aid and the destination hearing aid have different typ... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 13 | passed
Systems | passed | MASTER-22544 - The fitting software shall display the acoustic options available for the destin... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 14 | passed
Systems | passed | MASTER-22545 - The fitting software shall provide a method to change the acoustic option for th... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 15 | passed
Systems | passed | MASTER-22547 - For each enabled memory location, the fitting software shall use the gain/freque... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 16 | passed
Systems | passed | MASTER-22549 - When the original hearing aid and the destination hearing aid have different typ... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 17 | passed
Systems | passed | MASTER-22550 - When the original hearing aid and the destination hearing aid have different typ... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 18 | passed
Systems | passed | MASTER-22551 - When the original hearing aid has no means of volume control and the destination... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 19 | passed
Systems | passed | MASTER-22552 - If it is not possible to match the frequency response of the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 20 | passed
Systems | passed | MASTER-22553 - When neither the original hearing aid nor the destination hearing aid has a mean... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 21 | passed
Systems | passed | MASTER-22554 - When the original hearing aid has a means of volume control and the destination... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 22 | passed
Systems | passed | MASTER-22555 - The fitting software shall transfer the settings for the Memory Reset feature fr... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 313 | passed
Systems | passed | MASTER-22556 - When the original hearing aid does not have Smart VC capability and the destinat... | MASTER-28190 - SV195 - Smart Volume Control | Requirements Cleanup 4 - Polarion Cleanup Round 4 | Executed | 320 | passed
[5 empty] | Executed | 326 | passed
Systems | passed | MASTER-22559 - When the original hearing aid has 4 memory locations and the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 23 | passed
Systems | passed | MASTER-22560 - When the original hearing aid has 4 memory locations, the destination hearing ai... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 24 | passed
Systems | passed | MASTER-22561 - When the original hearing aid has 4 memory locations and the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 25 | passed
Systems | passed | MASTER-22563 - When the original hearing aid has 2 memory locations and the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 26 | passed
Systems | passed | MASTER-22564 - When the original hearing aid has 2 memory locations, the destination hearing ai... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 27 | passed
Systems | passed | MASTER-22565 - When the original hearing aid has 2 memory locations, the destination hearing ai... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 28 | passed
Systems | passed | MASTER-22567 - When the original hearing aid has 1 memory location and the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 29 | passed
Systems | passed | MASTER-22569 - When the original hearing aid has 1 memory location and the destination hearing... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 30 | passed
Systems | passed | MASTER-22571 - When the original hearing aid has a memory environment that is available on the... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 31 | passed
Systems | passed | MASTER-22572 - When the original hearing aid has a non-telephone memory environment that is not... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 32 | passed
Systems | passed | MASTER-22573 - When the original hearing aid has a telephone memory environment that is not ava... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 33 | passed
Systems | passed | MASTER-22574 - When the original hearing has a memory location that is disabled, the fitting so... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 34 | passed
Systems | passed | MASTER-22576 - When the original hearing aid and destination hearing aid both have 2 memory loc... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 35 | passed
Systems | passed | MASTER-22577 - When the original hearing aid and destination hearing aid both have 2 memory loc... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 36 | passed
Systems | passed | MASTER-22578 - When the original hearing aid and destination hearing aid both have 2 memory loc... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 37 | passed
Systems | passed | MASTER-22580 - The fitting software shall provide a method for wirelessly pairing the destinati... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 38 | passed
Systems | passed | MASTER-22581 - If conflicts are created due to the transfer of wireless settings to the destina... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 39 | passed
Systems | passed | MASTER-51240 - When both WiCROS Partners or only a WiCROS Partner hearing aid is connected, and... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 115 | passed
[5 empty] | Executed | 117 | passed
Systems | passed | MASTER-51241 - When both WiCROS Partners or only a WiCROS Partner hearing aid is connected, and... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 108 | passed
[5 empty] | Executed | 110 | passed
Systems | passed | MASTER-51238 - When binaural non-WiCROS hearing aids are connected, and a Swap Fit using a WiCR... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 102 | passed
Systems | passed | MASTER-51239 - When one non-WiCROS hearing aid is connected, and a Swap Fit using a WiCROS mona... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 105 | passed
Systems | passed | MASTER-51242 - When both WiCROS Partners or only a WiCROS Partner hearing aid is connected, and... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 122 | passed
[5 empty] | Executed | 124 | passed
Systems | passed | MASTER-51243 - When both WiCROS Partners or only a WiCROS Partner hearing aid is connected, and... | MASTER-53148 - SV245 - NFMI WiCROS - Indicators&AudioSync&Others | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 129 | passed
[5 empty] | Executed | 131 | passed
Systems | passed | MASTER-22584 - When the original hearing aid and the destination hearing aid both have Auto Acc... | MASTER-28182 - SV128 - Swap Fit | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 68 | passed
[5 empty] | Executed | 69 | passed
Systems | passed | MASTER-22585 - When the destination hearing aid has Auto Acclimatization capability, and the fe... | MASTER-28182 - SV128 - Swap Fit | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 76 | passed
[5 empty] | Executed | 77 | passed
Systems | passed | MASTER-22586 - When the original hearing aid has Auto Acclimatization capability but the destin... | MASTER-28182 - SV128 - Swap Fit | Amplifon Inspire - Testing build 24.0.10192.0, Suite 25.1.5057 | Executed | 85 | passed
[5 empty] | Executed | 86 | passed
Systems | passed | MASTER-22593 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 35 | passed
Systems | passed | MASTER-22594 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 40 | passed
Systems | passed | MASTER-22595 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 43 | passed
Systems | passed | MASTER-22596 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 46 | passed
Systems | passed | MASTER-22597 - When the following conditions are met: * The source/original hearing aid does no... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 53 | passed
Systems | passed | MASTER-22598 - When the following conditions are met: * The source/original hearing aid does no... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 58 | passed
Systems | passed | MASTER-21694 - When the following conditions are met: * The source/original hearing aid does no... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 55 | passed
[5 empty] | Executed | 58 | passed
Systems | passed | MASTER-21695 - When the following conditions are met: * The source/original hearing aid does no... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 61 | passed
Systems | passed | MASTER-21696 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 37 | passed
Systems | passed | MASTER-22600 - When the source/original hearing aid has Improved Indicator capability but the d... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-22601 - When the following conditions are met: * The source/original hearing aid does no... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 54 | passed
Systems | passed | MASTER-51771 - When the source/original hearing aid does not support 2.4 GHz accessory streamin... | MASTER-28189 - SV186E - Swapfit functionality with Improved Indicators | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 58 | passed
Systems | passed | MASTER-22603 - When the original hearing aid and the destination hearing aid both have HF Direc... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 103 | passed
Systems | passed | MASTER-22604 - When the destination hearing aid has HF Directionality capability but the origin... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 100 | passed
Systems | passed | MASTER-22605 - When the original hearing aid has HF directionality capability but the destinati... | MASTER-28192 - SV198 - High Frequency Directionality/Immersion Directionality | IP6 - General IP6 Testing | Executed | 106 | passed
Systems | passed | MASTER-22608 - When the destination hearing aid supports the Power Off user control function bu... | MASTER-34462 - SV204 - Rechargeable HA Feature testing | IP9 - General IP9 Testing & Regression | Executed | 49 | passed
[5 empty] | Executed | 54 | passed
Systems | passed | MASTER-22609 - When the original hearing aid Push and Hold gesture is assigned the Power Off us... | MASTER-34462 - SV204 - Rechargeable HA Feature testing | IP9 - General IP9 Testing & Regression | Executed | 58 | passed
Systems | passed | MASTER-22611 - The fitting software shall not transfer data from the original hearing aid to th... | MASTER-39840 - SV207 - Subsystem Verification - SV128 - Swap Fit | Requirements Cleanup 1 - Polarion Cleanup Round 1 | Executed | 41 | passed
Systems | passed | MASTER-50438 - When the Cloud Fitting Session is saved, the fitting software shall upload to th... | MASTER-47175 - SV228 - Remote Programming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 57 | passed
[5 empty] | Executed | 63 | passed
Fitting Software | passed | UTICA-11927 - Additional restrictions exist for Amplifon hearing aids and Amplifon editions of... | UTICA-22730 - Additional restrictions exist for Amplifon hearing aids and Amplifon editions of... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21839 - When the Tinnitus Stimulus is enabled in at least one hearing aid memory, and no... | UTICA-23012 - When the Tinnitus Stimulus is enabled in at least one hearing aid memory, and no... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21840 - When one or more user control gestures are configured to control the Tinnitus St... | UTICA-23008 - When one or more user control gestures are configured to control the Tinnitus St... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21891 - When a Volume Control feature is enabled and a Best Fit is performed, the fittin... | UTICA-23000 - When a Volume Control feature is enabled and a Best Fit is performed, the fittin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21897 - When E2E Volume Control is enabled in a binaural fitting, the fitting software s... | UTICA-22996 - When E2E Volume Control is enabled in a binaural fitting, the fitting software s... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21905 - When the default Reserve Gain is not available, the fitting software shall set t... | UTICA-22993 - When the default Reserve Gain is not available, the fitting software shall set t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21976 - For the Flagship tier, for each music-adaptation memory, when Music Adaption is... | UTICA-22990 - For the Flagship tier, for each music-adaptation memory, when Music Adaption is... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21977 - For the High-end tier, for each music-adaptation memory, when Music Adaption is... | UTICA-22987 - For the High-end tier, for each music-adaptation memory, when Music Adaption is... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21978 - For the Mid-end tier, when Music Adaption is enabled, the fitting software shall... | UTICA-22984 - For the Mid-end tier, when Music Adaption is enabled, the fitting software shall... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21979 - For the Low-end and Entry-level tiers, the fitting software shall preclude Music... | UTICA-22981 - For the Low-end and Entry-level tiers, the fitting software shall preclude Music... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21988 - For the Flagship tier, for each machine-adaptation memory, when Machine Adaption... | UTICA-22979 - For the Flagship tier, for each machine-adaptation memory, when Machine Adaption... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21989 - For the Upper-end tier, for each machine-adaptation memory, when Machine Adaptio... | UTICA-23045 - For the Upper-end tier, for each machine-adaptation memory, when Machine Adaptio... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21990 - For the Low-end and Entry-level tiers, when Machine Adaption is enabled, the fit... | UTICA-23042 - For the Low-end and Entry-level tiers, when Machine Adaption is enabled, the fit... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21997 - For the Flagship tier, for each wind-adaptation memory, when Wind Adaptation is... | UTICA-23039 - For the Flagship tier, for each wind-adaptation memory, when Wind Adaptation is... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21998 - For the Upper-end tier, for each wind-adaptation memory, when Wind Adaptation is... | UTICA-23036 - For the Upper-end tier, for each wind-adaptation memory, when Wind Adaptation is... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21999 - For the Mid-end, Low-end and Entry-level tiers, when Wind Adaptation is enabled,... | UTICA-23033 - For the Mid-end, Low-end and Entry-level tiers, when Wind Adaptation is enabled,... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50258 - For each memory location, the fitting software shall have a method to enable or... | UTICA-23030 - For each memory location, the fitting software shall have a method to enable or... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50256 - For each memory location, the fitting software shall display the enabled/disable... | UTICA-23027 - For each memory location, the fitting software shall display the enabled/disable... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50257 - The fitting software shall preclude Monaural Noise Reduction from being enabled... | UTICA-23024 - The fitting software shall preclude Monaural Noise Reduction from being enabled... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-76083 - The fitting software shall support the available Monaural Noise Reduction settin... | UTICA-26184 - The fitting software shall support the available Monaural Noise Reduction settin... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-50262 - For each memory location on hearing aids with NFMI, the fitting software shall h... | UTICA-23021 - For each memory location on hearing aids with NFMI, the fitting software shall h... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50260 - In a binaural fitting, the fitting software shall enable Binaural Noise Reductio... | UTICA-23019 - In a binaural fitting, the fitting software shall enable Binaural Noise Reductio... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50261 - The fitting software shall preclude Binaural Noise Reduction from being enabled... | UTICA-22935 - The fitting software shall preclude Binaural Noise Reduction from being enabled... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50259 - The fitting software shall preclude enabling Binaural Noise Reduction in a memor... | UTICA-22932 - The fitting software shall preclude enabling Binaural Noise Reduction in a memor... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50269 - The fitting software shall preclude Binaural Noise Reduction from being enabled... | UTICA-22929 - The fitting software shall preclude Binaural Noise Reduction from being enabled... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51170 - In a WiCROS fitting session, the fitting software shall preclude Binaural Noise... | UTICA-22926 - In a WiCROS fitting session, the fitting software shall preclude Binaural Noise... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50265 - For each memory location, the fitting software shall display the enabled/disable... | UTICA-22912 - For each memory location, the fitting software shall display the enabled/disable... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50266 - The fitting software shall preclude enabling Binaural Noise Reduction in any mem... | UTICA-22909 - The fitting software shall preclude enabling Binaural Noise Reduction in any mem... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50263 - When enabling Frequency Translation, the fitting software shall disable Binaural... | UTICA-22906 - When enabling Frequency Translation, the fitting software shall disable Binaural... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50264 - When a Best Fit is performed, the fitting software shall enable Binaural Noise R... | UTICA-22903 - When a Best Fit is performed, the fitting software shall enable Binaural Noise R... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50267 - When entering a binaural fitting session, if Binaural Noise Reduction is enabled... | UTICA-22900 - When entering a binaural fitting session, if Binaural Noise Reduction is enabled... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50268 - When entering a binaural fitting session, if Binaural Noise Reduction is enabled... | UTICA-22898 - When entering a binaural fitting session, if Binaural Noise Reduction is enabled... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51883 - For memory locations 1, 2, 3 and 4, the fitting software shall have a method to... | UTICA-22976 - For memory locations 1, 2, 3 and 4, the fitting software shall have a method to... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51886 - For each memory location, the fitting software shall display the enabled/disable... | UTICA-22973 - For each memory location, the fitting software shall display the enabled/disable... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51887 - For each memory location in a Flagship tier hearing aid, when Transient Noise Re... | UTICA-22970 - For each memory location in a Flagship tier hearing aid, when Transient Noise Re... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51888 - For each memory location in a High-end tier hearing aid, when Transient Noise Re... | UTICA-22967 - For each memory location in a High-end tier hearing aid, when Transient Noise Re... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51889 - For each memory location in the Mid-end, Low-end and Entry-level tier hearing ai... | UTICA-26185 - For each memory location in the Mid-end, Low-end and Entry-level tier hearing ai... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-51890 - When a Best Fit is performed, the fitting software shall enable Transient Noise... | UTICA-22950 - When a Best Fit is performed, the fitting software shall enable Transient Noise... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22008 - For the Flagship, the High-end, Mid-end tiers, for each memory location, the fit... | UTICA-22947 - For the Flagship, the High-end, Mid-end tiers, for each memory location, the fit... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22009 - For the Low-end tier, for each memory location, the fitting software shall have... | UTICA-22944 - For the Low-end tier, for each memory location, the fitting software shall have... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22010 - For the Entry-level tier, for each memory location, the fitting software shall h... | UTICA-22941 - For the Entry-level tier, for each memory location, the fitting software shall h... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22015 - For the Flagship tier, the fitting software shall have a method to collectively... | UTICA-22939 - For the Flagship tier, the fitting software shall have a method to collectively... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22016 - For the High-end, Mid-end, Low-end, and Entry-level tiers, the fitting software... | UTICA-22864 - For the High-end, Mid-end, Low-end, and Entry-level tiers, the fitting software... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | UTICA-25681 - The fitting software shall support the available Quiet settings for each device... | UTICA-26193 - The fitting software shall support the available Quiet settings for each device... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-53015 - For each memory location configured with a memory environment other than Telecoi... | UTICA-22852 - For each memory location configured with a memory environment other than Telecoi... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53019 - The fitting software shall preclude FBC from being enabled in the Telecoil, Auto... | UTICA-22849 - The fitting software shall preclude FBC from being enabled in the Telecoil, Auto... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53020 - For each memory location, when FBC is enabled, the fitting software shall have a... | UTICA-22846 - For each memory location, when FBC is enabled, the fitting software shall have a... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53021 - When MSG data exists for the hearing aid and the gain margin is greater than or... | UTICA-22843 - When MSG data exists for the hearing aid and the gain margin is greater than or... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53022 - When configured with a memory environment other than Music, MSG data exists for... | UTICA-22840 - When configured with a memory environment other than Music, MSG data exists for... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53023 - When configured with the Music memory environment, the fitting software shall de... | UTICA-22838 - When configured with the Music memory environment, the fitting software shall de... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53024 - The fitting software shall update the FBC default settings, if needed, when the... | UTICA-22894 - The fitting software shall update the FBC default settings, if needed, when the... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53025 - When MSG data exists for the hearing aid, the fitting software shall ensure that... | UTICA-22891 - When MSG data exists for the hearing aid, the fitting software shall ensure that... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53026 - When FBC initialization is successfully executed, the fitting software shall ens... | UTICA-22888 - When FBC initialization is successfully executed, the fitting software shall ens... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53027 - The fitting software shall allow the hearing aid programmed gain to be set to gr... | UTICA-22885 - The fitting software shall allow the hearing aid programmed gain to be set to gr... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53028 - The fitting software shall display the gain margin for each MECO channel. | UTICA-22882 - The fitting software shall display the gain margin for each MECO channel. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-53029 - When hearing aid programmed gain has been set to greater than the MSG, the fitti... | UTICA-22879 - When hearing aid programmed gain has been set to greater than the MSG, the fitti... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22191 - The fitting software shall configure the Auto-Streaming memory to the Stream Boo... | UTICA-22876 - The fitting software shall configure the Auto-Streaming memory to the Stream Boo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22224 - The fitting software balance control shall be able to adjust the WiCROS Partner... | UTICA-22873 - The fitting software balance control shall be able to adjust the WiCROS Partner... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22225 - Each adjustment of the balance control shall result in equal but opposite change... | UTICA-22870 - Each adjustment of the balance control shall result in equal but opposite change... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22274 - When a fitting session is started with a hearing aid that has not been branded b... | UTICA-22868 - When a fitting session is started with a hearing aid that has not been branded b... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55724 - When in a fitting session with a Demo hearing aid, the fitting software shall pr... | UTICA-22816 - When in a fitting session with a Demo hearing aid, the fitting software shall pr... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55726 - When in a fitting session with two Demo hearing aids, the fitting software shall... | UTICA-22814 - When in a fitting session with two Demo hearing aids, the fitting software shall... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55727 - When updating the product tier in a Demo hearing aid, the fitting software shall... | UTICA-22812 - When updating the product tier in a Demo hearing aid, the fitting software shall... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55728 - When updating the product tier in a Demo hearing aid, the fitting software shall... | UTICA-22810 - When updating the product tier in a Demo hearing aid, the fitting software shall... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55729 - When updating the product tier in a Demo hearing aid and a setting can't be main... | UTICA-22808 - When updating the product tier in a Demo hearing aid and a setting can't be main... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | UTICA-24748 - Prior to entering a fitting session, upon detection of a hearing aid powered by... | UTICA-26194 - Prior to entering a fitting session, upon detection of a hearing aid powered by... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | UTICA-24749 - Prior to entering a fitting session, upon detection and selection of a hearing a... | UTICA-26195 - Prior to entering a fitting session, upon detection and selection of a hearing a... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-55171 - When a hearing aid powered by a rechargeable lithium ion battery is in a fitting... | UTICA-22802 - When a hearing aid powered by a rechargeable lithium ion battery is in a fitting... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22282 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | UTICA-22800 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22283 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | UTICA-22798 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22284 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | UTICA-22820 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22285 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | UTICA-22835 - When Frequency Translation is enabled, the fitting software shall utilize the pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22332 - The Tinnitus Therapy feature shall be designed with a common programming interfa... | UTICA-22833 - The Tinnitus Therapy feature shall be designed with a common programming interfa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22333 - The fitting software shall provide the White Noise response of the Tinnitus Ther... | UTICA-22831 - The fitting software shall provide the White Noise response of the Tinnitus Ther... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64359 - The fitting software shall have a method to initiate one of the following respon... | UTICA-22829 - The fitting software shall have a method to initiate one of the following respon... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64686 - When the Tinnitus Therapy feature is enabled, the fitting software shall store t... | UTICA-22827 - When the Tinnitus Therapy feature is enabled, the fitting software shall store t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22335 - For each memory location in the hearing aid that is not configured as a Telephon... | UTICA-22825 - For each memory location in the hearing aid that is not configured as a Telephon... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21367 - For each memory location in the hearing aid that is not configured as a Telephon... | UTICA-22823 - For each memory location in the hearing aid that is not configured as a Telephon... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22336 - In a binaural fitting, the fitting software shall have a method to enable Tinnit... | UTICA-22821 - In a binaural fitting, the fitting software shall have a method to enable Tinnit... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22337 - For each memory location in the hearing aid that has Tinnitus Therapy enabled, t... | UTICA-22818 - For each memory location in the hearing aid that has Tinnitus Therapy enabled, t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22338 - The fitting software shall display the enabled/disabled status of the Tinnitus T... | UTICA-22749 - The fitting software shall display the enabled/disabled status of the Tinnitus T... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22339 - When a Best Fit of the hearing aid frequency response is performed, the fitting... | UTICA-22747 - When a Best Fit of the hearing aid frequency response is performed, the fitting... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22340 - The fitting software shall allow Tinnitus Therapy to be enabled with each of the... | UTICA-22745 - The fitting software shall allow Tinnitus Therapy to be enabled with each of the... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22342 - The fitting software shall have a method to adjust and set the Stimulus Level Co... | UTICA-22743 - The fitting software shall have a method to adjust and set the Stimulus Level Co... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21375 - For custom style products, the fitting software shall have a method to adjust an... | UTICA-22741 - For custom style products, the fitting software shall have a method to adjust an... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22344 - The fitting software shall have a method to adjust and program the Tinnitus Ther... | UTICA-22739 - The fitting software shall have a method to adjust and program the Tinnitus Ther... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22345 - The fitting software shall have a method to enable and disable the microphone du... | UTICA-22737 - The fitting software shall have a method to enable and disable the microphone du... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22346 - For a Stimulus Level Control step size of 2 dB, the fitting software shall provi... | UTICA-22755 - For a Stimulus Level Control step size of 2 dB, the fitting software shall provi... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22347 - For a Stimulus Level Control step size of 2 dB, the fitting software shall provi... | UTICA-22753 - For a Stimulus Level Control step size of 2 dB, the fitting software shall provi... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22348 - For a Stimulus Level Control step size of 4 dB, the fitting software shall provi... | UTICA-22752 - For a Stimulus Level Control step size of 4 dB, the fitting software shall provi... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22349 - For a Stimulus Level Control step size of 4 dB, the fitting software shall provi... | UTICA-22771 - For a Stimulus Level Control step size of 4 dB, the fitting software shall provi... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22350 - When E2E Tinnitus Therapy Stimulus Level Control is enabled in a binaural fittin... | UTICA-22769 - When E2E Tinnitus Therapy Stimulus Level Control is enabled in a binaural fittin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22356 - The fitting software shall have a method to adjust the overall Tinnitus Therapy... | UTICA-22767 - The fitting software shall have a method to adjust the overall Tinnitus Therapy... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64360 - The fitting software shall have a method to Mute and Unmute the hearing aid micr... | UTICA-22765 - The fitting software shall have a method to Mute and Unmute the hearing aid micr... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22358 - The fitting software shall display the frequency response and level of the Tinni... | UTICA-22763 - The fitting software shall display the frequency response and level of the Tinni... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21397 - The fitting software shall have a method to perform a Target Match of the Tinnit... | UTICA-22761 - The fitting software shall have a method to perform a Target Match of the Tinnit... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64362 - When the White Noise response is initiated, the fitting software shall utilize t... | UTICA-22759 - When the White Noise response is initiated, the fitting software shall utilize t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64364 - When the Audiogram-Shaped Stimulus response is initiated, the fitting software s... | UTICA-22757 - When the Audiogram-Shaped Stimulus response is initiated, the fitting software s... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64366 - The fitting software shall display the value of each band, in dB SPL, set by the... | UTICA-22774 - The fitting software shall display the value of each band, in dB SPL, set by the... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64690 - When Tinnitus Stimulus Customization response is selected, the fitting software... | UTICA-22773 - When Tinnitus Stimulus Customization response is selected, the fitting software... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64688 - During the MDL or MML test, the fitting software shall have a method to navigate... | UTICA-22704 - During the MDL or MML test, the fitting software shall have a method to navigate... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64693 - When Tinnitus Stimulus Customization is selected, the fitting software shall ope... | UTICA-22698 - When Tinnitus Stimulus Customization is selected, the fitting software shall ope... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64696 - Upon opening the MML test, the fitting software shall have a method to test and... | UTICA-22695 - Upon opening the MML test, the fitting software shall have a method to test and... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22362 - For each memory location in the hearing aid that has Tinnitus Therapy enabled, t... | UTICA-22692 - For each memory location in the hearing aid that has Tinnitus Therapy enabled, t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22363 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | UTICA-22689 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22364 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | UTICA-22686 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22365 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | UTICA-22683 - For each memory location in the hearing aid that has Tinnitus Therapy Stimulus f... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22371 - The fitting software shall utilize the same offsets from Level 3 (final gain set... | UTICA-22680 - The fitting software shall utilize the same offsets from Level 3 (final gain set... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22463 - The fitting software shall provide the Experience Manager feature for all hearin... | UTICA-22677 - The fitting software shall provide the Experience Manager feature for all hearin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22474 - When the fitting software is invoked with an audiogram that does not have a comp... | UTICA-22702 - When the fitting software is invoked with an audiogram that does not have a comp... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22497 - When the current experience level is set, the fitting software shall set the pre... | UTICA-22734 - When the current experience level is set, the fitting software shall set the pre... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22498 - When a hearing aid that has been previously fit enters a fitting session, and ei... | UTICA-22731 - When a hearing aid that has been previously fit enters a fitting session, and ei... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21585 - If a Swap Fit is performed to a pre-Milan destination hearing aid configured wit... | UTICA-22727 - If a Swap Fit is performed to a pre-Milan destination hearing aid configured wit... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22517 - When a Swap Fit from a Milan/Greater saved fitting session is performed to a pre... | UTICA-22724 - When a Swap Fit from a Milan/Greater saved fitting session is performed to a pre... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22518 - When a Swap Fit from a Pre-Milan saved fitting session is performed to a Milan/G... | UTICA-22721 - When a Swap Fit from a Pre-Milan saved fitting session is performed to a Milan/G... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21588 - When a Swap Fit from a Pre-Milan saved fitting session is performed to a Pre-Mil... | UTICA-22718 - When a Swap Fit from a Pre-Milan saved fitting session is performed to a Pre-Mil... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-22522 - When a session connect to a Pre-Milan saved fitting session with a Milan/Greater... | UTICA-22715 - When a session connect to a Pre-Milan saved fitting session with a Milan/Greater... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-21593 - When a session connect to a Pre-Milan saved fitting session with a Pre-Milan hea... | UTICA-22712 - When a session connect to a Pre-Milan saved fitting session with a Pre-Milan hea... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51319 - The fitting software shall have a method to enable and disable Fall Detection. | UTICA-22709 - The fitting software shall have a method to enable and disable Fall Detection. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-56919 - When Manual Alert feature-function is assigned to the Push and Hold user control... | UTICA-22707 - When Manual Alert feature-function is assigned to the Push and Hold user control... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51320 - When Fall Detection is enabled, the fitting software shall have a method to set... | UTICA-23015 - When Fall Detection is enabled, the fitting software shall have a method to set... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51326 - The fitting software shall have a method to clear the fall event data in the hea... | UTICA-23013 - The fitting software shall have a method to clear the fall event data in the hea... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51351 - The fitting software shall have a method to clear the manual alert data in the h... | UTICA-23010 - The fitting software shall have a method to clear the manual alert data in the h... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-51325 - If at least one fall event or manual alert has been logged in the hearing aid da... | UTICA-23002 - If at least one fall event or manual alert has been logged in the hearing aid da... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-56579 - When the binaural option is selected, the fitting software shall command one hea... | UTICA-22998 - When the binaural option is selected, the fitting software shall command one hea... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-57690 - .The fitting software shall have a method to disable and enable Auto Sleep. | UTICA-26186 - .The fitting software shall have a method to disable and enable Auto Sleep. | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | UTICA-8389 - The fitting software shall configure the On-Demand Adaptive Tuning feature envir... | UTICA-22992 - The fitting software shall configure the On-Demand Adaptive Tuning feature envir... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-60896 - The fitting software shall read the data log information from the hearing aid an... | UTICA-22989 - The fitting software shall read the data log information from the hearing aid an... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-60900 - The fitting software shall provide a method to enable and disable data logging i... | UTICA-22986 - The fitting software shall provide a method to enable and disable data logging i... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63139 - The software shall use the Noah API to connect to the target IMC protocol 2 real... | UTICA-22983 - The software shall use the Noah API to connect to the target IMC protocol 2 real... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63142 - The fitting software shall support aided real ear measurements at soft, moderate... | UTICA-26187 - The fitting software shall support aided real ear measurements at soft, moderate... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-63145 - AutoREM testing shall only be performed using normal memory. | UTICA-23046 - AutoREM testing shall only be performed using normal memory. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63144 - The AutoREM feature shall only support customers using Noah 4.5.1 or higher. | UTICA-23044 - The AutoREM feature shall only support customers using Noah 4.5.1 or higher. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63147 - The AutoREM feature shall support all NP5 and newer hearing aids. | UTICA-23041 - The AutoREM feature shall support all NP5 and newer hearing aids. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63146 - The AutoREM feature shall support the following third-party REM hardware and sof... | UTICA-26188 - The AutoREM feature shall support the following third-party REM hardware and sof... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-63151 - The AutoREM feature shall guide the user through the probe calibration procedure... | UTICA-23035 - The AutoREM feature shall guide the user through the probe calibration procedure... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63148 - The AutoREM feature shall explicitly show when the calibration is successful or... | UTICA-23032 - The AutoREM feature shall explicitly show when the calibration is successful or... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63149 - The AutoREM feature shall clearly indicate if the probe is outside of its allowe... | UTICA-23029 - The AutoREM feature shall clearly indicate if the probe is outside of its allowe... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63152 - The AutoREM feature shall inform the user if no supported measurement tool is co... | UTICA-23026 - The AutoREM feature shall inform the user if no supported measurement tool is co... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63153 - The AutoREM feature shall inform the user if the third-party REM equipment becom... | UTICA-23023 - The AutoREM feature shall inform the user if the third-party REM equipment becom... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63157 - The AutoREM feature shall provide on-screen user instructions at each step of th... | UTICA-22937 - The AutoREM feature shall provide on-screen user instructions at each step of th... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63160 - The user shall be warned that the wizard may alter fitting adjustments made prio... | UTICA-22936 - The user shall be warned that the wizard may alter fitting adjustments made prio... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63165 - The following AutoREM data shall be serialized and stored in the Inspire data wa... | UTICA-22934 - The following AutoREM data shall be serialized and stored in the Inspire data wa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63169 - Non-calibration AutoREM aided measurement graphs shall include: pre-adjustment S... | UTICA-22931 - Non-calibration AutoREM aided measurement graphs shall include: pre-adjustment S... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63170 - All AutoREM measurement graphs shall be clearly labeled with the ear they repres... | UTICA-22928 - All AutoREM measurement graphs shall be clearly labeled with the ear they repres... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63167 - All AutoREM graphs shall be distinguished between through the use of line types... | UTICA-22914 - All AutoREM graphs shall be distinguished between through the use of line types... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63168 - After measuring the real ear response, the AutoREM feature shall graph the Measu... | UTICA-22911 - After measuring the real ear response, the AutoREM feature shall graph the Measu... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-63173 - A user choosing to cancel out of the measurement process and having data that wo... | UTICA-22908 - A user choosing to cancel out of the measurement process and having data that wo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50306 - The cloud remote programming application shall have system tray functionality th... | UTICA-22905 - The cloud remote programming application shall have system tray functionality th... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50309 - Upon launching the cloud remote programming application, and when the cloud remo... | UTICA-22902 - Upon launching the cloud remote programming application, and when the cloud remo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50308 - When the cloud remote programming application loses connection with the cloud, t... | UTICA-22978 - When the cloud remote programming application loses connection with the cloud, t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50355 - The fitting software shall download or upload data from the cloud via the cloud... | UTICA-22977 - The fitting software shall download or upload data from the cloud via the cloud... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50312 - The cloud remote programming application shall have a method to create a Profess... | UTICA-22975 - The cloud remote programming application shall have a method to create a Profess... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50311 - When logged into a Professional-Admin cloud account, the cloud remote programmin... | UTICA-22972 - When logged into a Professional-Admin cloud account, the cloud remote programmin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50316 - The content of a Professional-Admin cloud account shall contain at least the fol... | UTICA-22969 - The content of a Professional-Admin cloud account shall contain at least the fol... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50314 - The cloud remote programming application shall have a method to enter the conten... | UTICA-22966 - The cloud remote programming application shall have a method to enter the conten... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50315 - When logged into a Professional-Admin cloud account, the cloud remote programmin... | UTICA-22952 - When logged into a Professional-Admin cloud account, the cloud remote programmin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50389 - If a Professional-User account was not verified, the cloud remote programming ap... | UTICA-22949 - If a Professional-User account was not verified, the cloud remote programming ap... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50387 - When logged in as a Professional-Admin, the cloud remote programming application... | UTICA-22946 - When logged in as a Professional-Admin, the cloud remote programming application... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50388 - If a Professional-User account is deleted, the cloud remote programming applicat... | UTICA-22943 - If a Professional-User account is deleted, the cloud remote programming applicat... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50318 - When logged into a Professional-Admin cloud account, and a Professional-User gro... | UTICA-22866 - When logged into a Professional-Admin cloud account, and a Professional-User gro... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50392 - When Professional-User is added to a Professional-User group, the cloud remote p... | UTICA-22865 - When Professional-User is added to a Professional-User group, the cloud remote p... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50320 - The content of the Professional-User cloud account shall contain at least the fo... | UTICA-22863 - The content of the Professional-User cloud account shall contain at least the fo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50390 - When logged in as a Professional-User who doesn't belong to any Professional-Use... | UTICA-22860 - When logged in as a Professional-User who doesn't belong to any Professional-Use... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50359 - When creating a cloud account, the cloud remote programming application shall ve... | UTICA-22857 - When creating a cloud account, the cloud remote programming application shall ve... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50360 - If the cloud account username is not an email address, the cloud remote programm... | UTICA-22854 - If the cloud account username is not an email address, the cloud remote programm... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50358 - If a cloud account using that username already exists, the cloud remote programm... | UTICA-22851 - If a cloud account using that username already exists, the cloud remote programm... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50365 - The cloud remote programming application shall verify the cloud account password... | UTICA-22848 - The cloud remote programming application shall verify the cloud account password... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50363 - If the cloud account password does not meet the password restrictions, the cloud... | UTICA-22845 - If the cloud account password does not meet the password restrictions, the cloud... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50364 - If the cloud account password meets the password restrictions, the cloud remote... | UTICA-22842 - If the cloud account password meets the password restrictions, the cloud remote... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50368 - When a request to create a cloud account is submitted, the cloud remote programm... | UTICA-22896 - When a request to create a cloud account is submitted, the cloud remote programm... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50369 - Upon receiving a notification to verify the Professional-Admin or Professional-U... | UTICA-22895 - Upon receiving a notification to verify the Professional-Admin or Professional-U... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50366 - If the cloud account username is not verified, the cloud remote programming appl... | UTICA-22893 - If the cloud account username is not verified, the cloud remote programming appl... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50367 - If the cloud account username is not verified in 7 days, the cloud remote progra... | UTICA-22890 - If the cloud account username is not verified in 7 days, the cloud remote progra... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50370 - If the cloud account username is verified, the cloud remote programming applicat... | UTICA-22887 - If the cloud account username is verified, the cloud remote programming applicat... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50371 - When the Professional-User content is entered by the Professional-Admin, the clo... | UTICA-22884 - When the Professional-User content is entered by the Professional-Admin, the clo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50324 - The cloud remote programming application shall have a method to: * Log into a cl... | UTICA-22881 - The cloud remote programming application shall have a method to: * Log into a cl... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50322 - When logging into a cloud account, the cloud remote programming application shal... | UTICA-22878 - When logging into a cloud account, the cloud remote programming application shal... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50378 - When logging into a cloud account, If the entered cloud account username or pass... | UTICA-22875 - When logging into a cloud account, If the entered cloud account username or pass... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50379 - When logging into a cloud account, If the cloud account username and cloud accou... | UTICA-22872 - When logging into a cloud account, If the cloud account username and cloud accou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50376 - If the cloud remote programming application loses internet connection while logg... | UTICA-22782 - If the cloud remote programming application loses internet connection while logg... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50377 - When logged into a cloud account, the cloud remote programming application shall... | UTICA-22781 - When logged into a cloud account, the cloud remote programming application shall... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50380 - After opening the cloud remote programming application, and if the previous clou... | UTICA-22780 - After opening the cloud remote programming application, and if the previous clou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50385 - When the cloud remote programming application is not logged into the cloud accou... | UTICA-22779 - When the cloud remote programming application is not logged into the cloud accou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50383 - Upon exiting the cloud remote programming application, the cloud remote programm... | UTICA-22778 - Upon exiting the cloud remote programming application, the cloud remote programm... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50384 - Upon logging out of the cloud account, the cloud remote programming application... | UTICA-22777 - Upon logging out of the cloud account, the cloud remote programming application... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50323 - The cloud remote programming application shall have a method to shut down via th... | UTICA-22786 - The cloud remote programming application shall have a method to shut down via th... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50326 - When logged out of the cloud, the cloud remote programming application shall hav... | UTICA-22785 - When logged out of the cloud, the cloud remote programming application shall hav... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50373 - Upon receiving a cloud account password reset request, the cloud remote programm... | UTICA-22784 - Upon receiving a cloud account password reset request, the cloud remote programm... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50374 - Upon receiving a notification for a password reset request, the cloud shall dire... | UTICA-22783 - Upon receiving a notification for a password reset request, the cloud shall dire... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-67082 - The cloud remote programming application shall have a method to change a cloud a... | UTICA-22793 - The cloud remote programming application shall have a method to change a cloud a... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-67083 - When logging into a cloud account and the password expiration time has been exce... | UTICA-22792 - When logging into a cloud account and the password expiration time has been exce... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-67080 - When the password expiration time becomes exceeded while logged into a cloud acc... | UTICA-22791 - When the password expiration time becomes exceeded while logged into a cloud acc... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-67081 - When a cloud account password is changed, the cloud remote programming applicati... | UTICA-22790 - When a cloud account password is changed, the cloud remote programming applicati... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50305 - The fitting software shall have a method to enable or disable Remote Programming... | UTICA-22789 - The fitting software shall have a method to enable or disable Remote Programming... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55257 - In a binaural fitting session, the fitting software shall enable or disable Remo... | UTICA-22788 - In a binaural fitting session, the fitting software shall enable or disable Remo... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50328 - The fitting software shall have a method to generate a Remote Programming Code. | UTICA-22787 - The fitting software shall have a method to generate a Remote Programming Code. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50396 - When a Remote Programming Code is requested, the cloud remote programming applic... | UTICA-22796 - When a Remote Programming Code is requested, the cloud remote programming applic... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50395 - The Remote Programming Code shall be unique across the entire cloud. | UTICA-22795 - The Remote Programming Code shall be unique across the entire cloud. | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50398 - If a patient is selected in NOAH version 4.3 or after and a Remote Programming C... | UTICA-22794 - If a patient is selected in NOAH version 4.3 or after and a Remote Programming C... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50401 - If a Remote Programming Code is requested, and the office system version is prio... | UTICA-22751 - If a Remote Programming Code is requested, and the office system version is prio... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50400 - A General Remote Programming Code shall be used by the first patient who enters... | UTICA-22750 - A General Remote Programming Code shall be used by the first patient who enters... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50330 - The cloud remote programming application shall poll the cloud for Help Requests... | UTICA-22748 - The cloud remote programming application shall poll the cloud for Help Requests... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50331 - The cloud remote programming application shall have a method to manually poll th... | UTICA-22746 - The cloud remote programming application shall have a method to manually poll th... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50405 - When the cloud remote programming application polls the cloud, the cloud remote... | UTICA-22744 - When the cloud remote programming application polls the cloud, the cloud remote... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55094 - When the option to receive an email notification when a Help Request is submitte... | UTICA-22742 - When the option to receive an email notification when a Help Request is submitte... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50403 - If the cloud polling failed within 60 seconds of initiating the poll, the cloud... | UTICA-22740 - If the cloud polling failed within 60 seconds of initiating the poll, the cloud... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50404 - When the cloud receives a request from the cloud remote programming application... | UTICA-22738 - When the cloud receives a request from the cloud remote programming application... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64619 - When in the Synchronous Remote Programming View, the cloud remote programming ap... | UTICA-22756 - When in the Synchronous Remote Programming View, the cloud remote programming ap... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64621 - The cloud remote programming application shall have a method to view a list of t... | UTICA-22754 - The cloud remote programming application shall have a method to view a list of t... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64622 - The cloud remote programming application shall display the following for each pa... | UTICA-22776 - The cloud remote programming application shall display the following for each pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64625 - Upon logging into the cloud, the cloud remote programming application shall sort... | UTICA-22772 - Upon logging into the cloud, the cloud remote programming application shall sort... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64624 - The cloud remote programming application shall have a method to search the list... | UTICA-22770 - The cloud remote programming application shall have a method to search the list... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64626 - The cloud remote programming application shall have a method to view the history... | UTICA-22768 - The cloud remote programming application shall have a method to view the history... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64627 - The cloud remote programming application shall display the history of Synchronou... | UTICA-22766 - The cloud remote programming application shall display the history of Synchronou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64630 - The cloud remote programming application shall display the following for each en... | UTICA-22764 - The cloud remote programming application shall display the following for each en... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64629 - The cloud remote programming application shall have a method to sort entries in... | UTICA-22762 - The cloud remote programming application shall have a method to sort entries in... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64632 - The cloud remote programming application shall have a method to initiate a Synch... | UTICA-22760 - The cloud remote programming application shall have a method to initiate a Synch... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | UTICA-6760 - Upon attempting to initiate a Synchronous Remote Fitting Session and the fitting... | UTICA-22758 - Upon attempting to initiate a Synchronous Remote Fitting Session and the fitting... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64634 - When a Synchronous Remote Fitting Session is initiated and the patient is not wa... | UTICA-22775 - When a Synchronous Remote Fitting Session is initiated and the patient is not wa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64637 - When a patient is waiting to join a Synchronous Remote Fitting Session and the S... | UTICA-22706 - When a patient is waiting to join a Synchronous Remote Fitting Session and the S... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64636 - When a Synchronous Remote Fitting Session is initiated and the patient is waitin... | UTICA-22705 - When a Synchronous Remote Fitting Session is initiated and the patient is waitin... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64638 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | UTICA-22700 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64639 - When the fitting software starts a Synchronous Remote Fitting Session, the fitti... | UTICA-22697 - When the fitting software starts a Synchronous Remote Fitting Session, the fitti... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-75378 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | UTICA-26189 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-75377 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | UTICA-26190 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | UTICA-24666 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | UTICA-26191 - When the fitting software starts a Synchronous Remote Fitting Session for a sele... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Fitting Software | passed | MASTER-64642 - The cloud remote programming application shall have a method to enable and disab... | UTICA-22685 - The cloud remote programming application shall have a method to enable and disab... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64641 - The cloud remote programming application shall have a method to mute and unmute... | UTICA-22682 - The cloud remote programming application shall have a method to mute and unmute... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64644 - When the fitting software changes to a different memory location during the Sync... | UTICA-22679 - When the fitting software changes to a different memory location during the Sync... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64643 - During a Synchronous Remote Fitting Session, the fitting software shall have a m... | UTICA-22736 - During a Synchronous Remote Fitting Session, the fitting software shall have a m... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64646 - During a Synchronous Remote Fitting Session, the fitting software shall provide... | UTICA-22735 - During a Synchronous Remote Fitting Session, the fitting software shall provide... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64645 - For Memory 1, Memory 2, Memory 3, Memory 4 and Memory 6 (StreamBoost), the fitti... | UTICA-22733 - For Memory 1, Memory 2, Memory 3, Memory 4 and Memory 6 (StreamBoost), the fitti... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64648 - For Custom memories, the fitting software shall require adjustments to be upload... | UTICA-22729 - For Custom memories, the fitting software shall require adjustments to be upload... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64647 - The fitting software shall allow additional adjustments to be made while previou... | UTICA-22726 - The fitting software shall allow additional adjustments to be made while previou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64649 - During a Synchronous Remote Fitting Session, when the connection to the the clou... | UTICA-22723 - During a Synchronous Remote Fitting Session, when the connection to the the clou... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-66516 - When a Synchronous Remote Fitting Session connection to the cloud is lost and ca... | UTICA-22720 - When a Synchronous Remote Fitting Session connection to the cloud is lost and ca... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64650 - The fitting software shall have a method to end the Synchronous Remote Fitting S... | UTICA-22717 - The fitting software shall have a method to end the Synchronous Remote Fitting S... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64651 - When adjustments have been made that have not been uploaded to the cloud and end... | UTICA-22714 - When adjustments have been made that have not been uploaded to the cloud and end... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-64653 - When a Synchronous Remote Fitting Session is ended, the cloud remote programming... | UTICA-22711 - When a Synchronous Remote Fitting Session is ended, the cloud remote programming... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50350 - If the Professional-User or Professional-Admin cloud account is linked to the Pa... | UTICA-23017 - If the Professional-User or Professional-Admin cloud account is linked to the Pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-55258 - Upon starting a binaural Cloud Fitting Session, when the enabled/disabled state... | UTICA-23016 - Upon starting a binaural Cloud Fitting Session, when the enabled/disabled state... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50436 - If the Professional-Admin or Professional-User cloud account is linked to the Pa... | UTICA-23014 - If the Professional-Admin or Professional-User cloud account is linked to the Pa... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50437 - Upon starting a Cloud Fitting Session, the fitting software shall download from... | UTICA-23011 - Upon starting a Cloud Fitting Session, the fitting software shall download from... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50434 - During a Cloud Fitting Session, if there is a Help Request Response with a memor... | UTICA-23003 - During a Cloud Fitting Session, if there is a Help Request Response with a memor... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | MASTER-50438 - When the Cloud Fitting Session is saved, the fitting software shall upload to th... | UTICA-22999 - When the Cloud Fitting Session is saved, the fitting software shall upload to th... | Software - Systems - Verification of Systems req. | Executed | N/A | passed
Fitting Software | passed | UTICA-25736 - The fitting software shall contain a website link to its electronic Operations M... | UTICA-26192 - The fitting software shall contain a website link to its electronic Operations M... | Additional SyRS Requirements for Utica - LP | Executed | N/A | passed
Regulatory | passed | MASTER-42207 - The fitting software shall include a UDI that is displayed to the user in easily... | UTICA-24663 - Utica Inspire SyRS - Regulatory | Utica - Regulatory | Executed | 1 | passed
Regulatory | passed | MASTER-42208 - When the fitting software is provided via physical media, the media, the packagi... | UTICA-24663 - Utica Inspire SyRS - Regulatory | Utica - Regulatory | Executed | 2 | passed
Regulatory | passed | MASTER-42209 - The UDI shall incorporate a Device Identifier which is different for each brand... | UTICA-24663 - Utica Inspire SyRS - Regulatory | Utica - Regulatory | Executed | 3 | passed