## Sheet1
Utica Features on Paris Wired Customs HA Feature Metrics Report Grand Avenue Document ID: GQ-RPT-01040 Revision: 1.0
Description of data in this document: Function: This column designates which functional team is responsible for verifying the corresponding Systems requirement. Requirement Result: This column indicates the cumulative (across final test runs) verification result for the corresponding requirement. This is the cumulative result of the corresponding Test Resullt (Column H) data. Requirement: This column lists the requirement title as shown in Polarion and provides a hyperlink to the requirement in Polarion. Test Case: This column indicates which test case(s) contain steps that verify the requirement. Please note that Polarion includes all executed tests that reference the relevant requirement. Please also note that Polarion only includes verification results in this report from the most recently executed Test Run(s). For these reasons, some results included in this report may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Run: This column lists the test run(s) that contained the test case(s) (Column D) that verify the requirement. Please note that Polarion only includes verification results in this report from the most recently executed Test Run(s), which may or may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Result Status: This column indicates the execution status for each instance of the requirement in all executed test cases. Test Step: This column indicates the location of every instance of the requirement in the corresponding Test Case (Column D). Zero-step test cases (test cases that directly pass/fail an individual requirement) will indicate N/A in this column. Test Result: This column indicates the test result for every instance of the requirement in the corresponding Test Case (Column D). This column is the input for the Requirement Result (Column B) column.
Function | Requirement Result | Requirement | Test Case | Test Run | Test Result Status | Test Step | Test Result
Firmware | passed | UTICA-9576 - With the exception of MNR and OROCL, the CIC hearing aids shall include the same... | MASTER-74585 - Napa Requirement Test Procedure Collection (WiredOnlyFW on Custom Devices) | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 46 | passed
Firmware | passed | UTICA-9577 - With the exception of MNR and OROCL, the IIC hearing aid shall include the same... | MASTER-74585 - Napa Requirement Test Procedure Collection (WiredOnlyFW on Custom Devices) | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 46 | passed
Firmware | passed | MASTER-50242 - When Monaural Noise Reduction is enabled and Binaural Noise Reduction is not act... | MASTER-54864 - BNR / MNR Switching Logic Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-42125 - When more than one gain reduction is simultaneously requested by Machine adaptat... | MASTER-74216 - AEC Machine Noise with and without Wind in Different Microphone Modes (Telsa, MNR) | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-74217 - AEC Other Noise with and without Wind in Different Microphone Modes (Tesla, MNR) | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-54930 - E2E Machine Noise Adaptation | HA-Utica_5_12_2021 - Utica Hopper Regression Testing | Executed | 1 | passed
|  |  | MASTER-55003 - E2E Machine Noise Adaptation WiCROS Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 2 | passed
Firmware | passed | MASTER-53031 - When FBC is enabled, the hearing aid shall apply feedback cancellation to MECO c... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 1 | passed
[5 empty] | Executed | 5 | passed
Firmware | passed | MASTER-53034 - When FBC has been set to Adaptive - High Sensitivity, the hearing aid shall shif... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53035 - When FBC has been set to Adaptive - Low Sensitivity, the hearing aid shall shift... | MASTER-74721 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53036 - When FBC has been set to Static, the hearing aid shall not apply OPM to the hear... | MASTER-74721 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_5_24_2021 - Utica Tesla Regression Testing | Executed | 2 | passed
|  |  | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54306 - Adaptive FBC High/Low Sensitivity Test | HA-Utica_3_17_2021 - Utica Faraday Regression Testing | Executed | 1 | passed
Firmware | passed | MASTER-53038 - The hearing aid shall apply OPM in MECO channels 5 through 24 when the energy in... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53039 - The hearing aid shall apply OPM in MECO channels 6 through 24 when either of the... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Firmware | passed | MASTER-53040 - The hearing aid shall apply OPM in MECO channels 8 through 24 when the energy in... | MASTER-74566 - FBC DUT WiredOnlyFW Test (FBC_simple_Tesla.ste) | HA-Utica_6_25_2021 - Utica-Tesla Automated Testing | Executed | 1 | passed
|  |  | MASTER-54265 - Feedback Canceller (Automated - Digital Unit Test) | HA-Utica_5_13_2021 - Utica Automated Regression Testing | Executed | 5 | passed
Systems | passed | UTICA-9576 - With the exception of MNR and OROCL, the CIC hearing aids shall include the same... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 4 | passed
Systems | passed | UTICA-9577 - With the exception of MNR and OROCL, the IIC hearing aid shall include the same... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 5 | passed
Systems | passed | MASTER-50245 - When a stationary white noise field is present and when configured in the curren... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 54 | passed
Systems | passed | MASTER-50248 - When an 8-talker babble input is present and when configured in the current memo... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 75 | passed
Systems | passed | MASTER-50250 - When presented with MNR_BNR_Speech_Input_1 and when configured in the current me... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 87 | passed
Systems | passed | MASTER-50249 - When presented with a 70 dB SPL mixture of MNR_BNR_Speech_Input_1 and 8-talker b... | MASTER-46616 - SV226 - MNR Performance Test | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 99 | passed
Systems | passed | MASTER-51774 - When MNR is enabled and the input level is less than 45 dB SPL, the hearing aid... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 121 | passed
Systems | passed | MASTER-51778 - For the Music memory environment, when the input level is less than 45 dB SPL, t... | MASTER-46195 - SV223 - MNR - below TK | IP10 - Bug Fixes & Regression | Executed | 123 | passed
Systems | passed | MASTER-42125 - When more than one gain reduction is simultaneously requested by Machine adaptat... | MASTER-33350 - SV051 - AEC Interactions | IP7 - General IP7 Testing | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 27 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 41 | passed
[5 empty] | Executed | 44 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 50 | passed
[5 empty] | Executed | 62 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 66 | passed
Systems | passed | MASTER-42000 - For each MECO channel, the hearing aid shall apply OCL gain reduction based on t... | MASTER-28145 - AV045 - Wide Band and Sub Band OCL | Utica NW CIC and IIC - Non Wireless CIC and IIC (Tesla FW) Testing | Executed | 37 | passed
Systems | passed | UTICA-11528 - The hearing aid shall limit total output such that Total Harmonic Distortion is... | UTICA-24332 - Utica Test By Analysis | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 6 | passed
Systems | passed | MASTER-19947 - The hearing aid shall be programmable using all of the following wired programme... | MASTER-26354 - SV107 - Wired Programmers | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 11 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 23 | passed
[5 empty] | Executed | 28 | passed
[5 empty] | Executed | 35 | passed
[5 empty] | Executed | 40 | passed
[5 empty] | Executed | 47 | passed
[5 empty] | Executed | 52 | passed
[5 empty] | Executed | 59 | passed
[5 empty] | Executed | 64 | passed
Systems | passed | MASTER-19949 - The hearing aid shall be programmable using at least one of the following cables... | MASTER-26354 - SV107 - Wired Programmers | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 11 | passed
[5 empty] | Executed | 16 | passed
[5 empty] | Executed | 72 | passed
[5 empty] | Executed | 77 | passed
[5 empty] | Executed | 84 | passed
[5 empty] | Executed | 89 | passed
[5 empty] | Executed | 96 | passed
[5 empty] | Executed | 101 | passed
[5 empty] | Executed | 109 | passed
[5 empty] | Executed | 114 | passed
[5 empty] | Executed | 121 | passed
[5 empty] | Executed | 126 | passed