## Sheet1
Utica 2.4 GHz CIC 312 Zn-Air Metrics Report Grand Avenue Document ID: GQ-RPT-01044 Revision: 1.0
Description of data in this document: Function: This column designates which functional team is responsible for verifying the corresponding Systems requirement. Requirement Result: This column indicates the cumulative (across final test runs) verification result for the corresponding requirement. This is the cumulative result of the corresponding Test Resullt (Column H) data. Requirement: This column lists the requirement title as shown in Polarion and provides a hyperlink to the requirement in Polarion. Test Case: This column indicates which test case(s) contain steps that verify the requirement. Please note that Polarion includes all executed tests that reference the relevant requirement. Please also note that Polarion only includes verification results in this report from the most recently executed Test Run(s). For these reasons, some results included in this report may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Run: This column lists the test run(s) that contained the test case(s) (Column D) that verify the requirement. Please note that Polarion only includes verification results in this report from the most recently executed Test Run(s), which may or may not address the subject of this report (due to the fact that requirements may be shared between multiple devices/configurations). The results pertinent to the subject of this report can be found in Polarion's live database and can be demonstrated upon request. Test Result Status: This column indicates the execution status for each instance of the requirement in all executed test cases. Test Step: This column indicates the location of every instance of the requirement in the corresponding Test Case (Column D). Zero-step test cases (test cases that directly pass/fail an individual requirement) will indicate N/A in this column. Test Result: This column indicates the test result for every instance of the requirement in the corresponding Test Case (Column D). This column is the input for the Requirement Result (Column B) column.
Function | Requirement Result | Requirement | Test Case | Test Run | Test Result Status | Test Step | Test Result
Wireless | passed | MASTER-20000 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-48455 - The hearing aid shall be capable of communicating with compatible Starkey access... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20001 - Design Guideline: It is desired that the hearing aid be capable of communicating... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20002 - Design Guideline: With no line-of-sight from the BLE Central to the user's head... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-20011 - The hearing aid shall receive commands from the Remote Control at a distance of... | MASTER-51617 - Hearing Aid Communication Range | U30 - SyRS-WiRS Traceability | Executed | 2 | passed
[5 empty] | Executed | 2 | passed
Wireless | passed | MASTER-61163 - The hearing aid shall be capable of E2E communication for at least the percentag... | MASTER-73153 - E2E Hardware Performance - Custom CIC Hearing Aids | Hearing Aid RF Hardware - CIC | Executed | 3 | passed
Wireless | passed | MASTER-20027 - The hearing aid shall achieve a link availability of greater than or equal to 95... | MASTER-51619 - Hearing Aid Communication Reliability | Hearing Aid RF HW - P BTE | Executed | 2 | passed
Wireless | passed | MASTER-20034 - The hearing aid shall have an in-situ TRP of better than or equal to -26 dBm as... | MASTER-48355 - Radiated Power | Hearing Aid RF Hardware - CIC | Executed | 5 | passed
Wireless | passed | MASTER-20035 - The hearing aid shall have an in-situ TIS of better than or equal to -66 dBm at... | MASTER-48377 - Total Isotropic Sensitivity - LE 1M PHY | Hearing Aid RF Hardware - CIC | Executed | 6 | passed
Wireless | passed | MASTER-20049 - The hearing aid shall communicate wirelessly using Bluetooth Low Energy as defin... | MASTER-49133 - Bluetooth Low Energy Core Compliance | Hearing Aid RF Hardware - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20113 - The hearing aid shall maintain wireless communication when exposed to an RF fiel... | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 9 | passed
Wireless | passed | MASTER-20143 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20144 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 5 | passed
Wireless | passed | MASTER-20145 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 6 | passed
Wireless | passed | MASTER-20147 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 7 | passed
Wireless | passed | MASTER-20148 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | MASTER-48782 - United States Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 9 | passed
Wireless | passed | MASTER-20152 - The hearing aid shall comply with ERC/REC 70-03, Annex 3, frequency band A. | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 5 | passed
Wireless | passed | MASTER-20153 - The hearing aid shall comply with Commission Decision 2009/381/EC for operation... | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 6 | passed
Wireless | passed | MASTER-20154 - The hearing aid shall comply with ETSI EN 300 328 for operation within the band... | MASTER-48785 - European Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20158 - The hearing aid shall comply with the requirements of RSS-Gen that are applicabl... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20159 - The hearing aid shall comply with RSS-247 for operation within the band of 2400-... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 6 | passed
Wireless | passed | MASTER-20160 - The hearing aid shall comply with RSS-102 for operation within the band of 2400-... | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 9 | passed
Wireless | passed | MASTER-20161 - The hearing aid shall comply with RSP-100. | MASTER-48783 - Canada Wireless Regulatory Compliance | HA Wireless Regulatory - CIC | Executed | 8 | passed
Wireless | passed | MASTER-20164 - The hearing aid shall comply with the performance and methods of measurement of... | MASTER-48794 - Australia Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20167 - The hearing aid shall comply with KN 301 489-1. Note: Testing is performed using... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 6 | passed
Wireless | passed | MASTER-42289 - The hearing aid shall comply with KN 301 489-17. Note: Testing is performed usin... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 5 | passed
Wireless | passed | MASTER-42290 - The hearing aid shall comply with Korean standard KN 300 328 for operation withi... | MASTER-48792 - Korea Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20170 - The hearing aid shall comply with China Ministry of Industry and Information Tec... | MASTER-48791 - China Wireless Regulatory Conformance | Hearing Aid Wireless Regulatory - P BTE | Executed | 4 | passed
Wireless | passed | MASTER-20172 - The hearing aid shall comply with Japan Association of Radio Industries and Busi... | MASTER-48790 - Japan Wireless Regulatory Conformance | HA Wireless Regulatory - CIC | Executed | 4 | passed
Wireless | passed | MASTER-20174 - The hearing aid shall be qualified to the standards defined by the Bluetooth Spe... | MASTER-49133 - Bluetooth Low Energy Core Compliance | Hearing Aid RF Hardware - CIC | Executed | 4 | passed
Systems | passed | MASTER-17316 - The hearing aid shall be capable of electrical operation using a battery from al... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 21 | passed
[5 empty] | Executed | 26 | passed
[5 empty] | Executed | 74 | passed
Systems | passed | MASTER-48450 - When powered by a size 312 zinc air battery, the hearing aid in all matrix confi... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 24 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-53165 - When powered by a size 312 zinc air battery, it is desired that the hearing aid... | MASTER-16450 - SV113 - Battery Life With Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 25 | passed
[5 empty] | Executed | 35 | passed
Systems | passed | MASTER-19838 - The tinnitus masking device shall be capable of generating a broadband noise wit... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 41 | passed
Systems | passed | MASTER-19839 - The frequency response of the tinnitus masking device noise shall be adjustable... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19840 - The level of the tinnitus masking device noise shall be manually adjustable thro... | MASTER-49033 - AV172 - Mobile Application - Mobile Application and Controls | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 130 | passed
|  |  | MASTER-48991 - AV173 Mobile Application and Controls - Additional | IP9 - General IP9 Testing & Regression | Executed | 106 | passed
[5 empty] | Executed | 106 | passed
|  |  | MASTER-26723 - SV002C - User Control - Tinnitus | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 137 | passed
[5 empty] | Executed | 247 | passed
Systems | passed | MASTER-19841 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 36 | passed
Systems | passed | MASTER-19842 - For a tinnitus masking device that provides noise and hearing aid functionality,... | MASTER-42069 - SV140 - Tinnitus Manual Acoustic Measurements | Requirements Cleanup 2 - Polarion Cleanup Round 2 | Executed | 35 | passed
Systems | passed | MASTER-48451 - The hearing aid shall produce no audible artifacts that are deemed objectionable... | MASTER-16449 - SV001 - General Audio Artifact | PPBTE IP5 - PPBTE SyRS Requirements Run 2 | Executed | 55 | passed
Systems | passed | UTICA-2476 - If the requirement for no audible artifacts while programming can't be achieved,... | MASTER-16449 - SV001 - General Audio Artifact | PPBTE IP5 - PPBTE SyRS Requirements Run 2 | Executed | 56 | passed
Systems | passed | MASTER-19678 - When Best Fit for a Flat 60 hearing loss, the hearing aid shall meet the perform... | MASTER-34339 - AV085 - Output Noise Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 41 | passed
[5 empty] | Executed | 73 | passed
|  |  | MASTER-34340 - AV084 - Output Noise Measurement | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
[5 empty] | Executed | 84 | passed
Systems | passed | UTICA-25732 - The hearing aid output noise shall not exceed the output noise limits defined in... | MASTER-34340 - AV084 - Output Noise Measurement | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 67 | passed
Systems | passed | UTICA-25733 - The hearing aid output noise shall not exceed the output noise limits defined in... | MASTER-34339 - AV085 - Output Noise Streaming | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 41 | passed
[5 empty] | Executed | 73 | passed
Systems | passed | MASTER-20000 - The hearing aid shall be capable of communicating with a BLE Central and compati... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-48455 - The hearing aid shall be capable of communicating with compatible Starkey access... | MASTER-49135 - SV234 - MFA Use Cases | Model S IP 6-5 | Executed | 163 | passed
[5 empty] | Executed | 166 | passed
[5 empty] | Executed | 169 | passed
[5 empty] | Executed | 172 | passed
|  |  | MASTER-52199 - SV241- Quebec Remote Mic- Accessories | Model S IP 6-5 | Executed | 36 | passed
[5 empty] | Executed | 40 | passed
Systems | passed | MASTER-20001 - Design Guideline: It is desired that the hearing aid be capable of communicating... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 125 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
[5 empty] | Executed | 136 | passed
[5 empty] | Executed | 138 | passed
Systems | passed | MASTER-20002 - Design Guideline: With no line-of-sight from the BLE Central to the user's head... | MASTER-16451 - SV120 - Wireless Connection Stability | IP9 - General IP9 Testing & Regression | Executed | 108 | passed
[5 empty] | Executed | 112 | passed
[5 empty] | Executed | 117 | passed
[5 empty] | Executed | 119 | passed
[5 empty] | Executed | 130 | passed
[5 empty] | Executed | 131 | passed
[5 empty] | Executed | 133 | passed
Systems | passed | MASTER-20011 - The hearing aid shall receive commands from the Remote Control at a distance of... | MASTER-34405 - SV117 - Remote Control Performance | CIC DVT 1 Run 1 - CIC SyRS Requirements Run 1 | Executed | 37 | passed
[5 empty] | Executed | 65 | passed
[5 empty] | Executed | 94 | passed
Systems | passed | MASTER-20064 - The hearing aid shall be programmable using the Starkey 2.4 GHz Programmer. | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 8 | passed
Systems | passed | MASTER-20065 - The hearing aid shall be programmable using the NOAHLink Wireless Programmer. No... | MASTER-34255 - SV007 - Wireless Functionality with Inspire | Requirements Cleanup 3 - Polarion Cleanup Round 3 | Executed | 8 | passed
ElectroAcoustics | passed | MASTER-17131 - The hearing aid shall use the NP5 platform. | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-42251 - The hearing aid shall include a Bluetooth radio. | UTICA-2760 - Verify by Inspection | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-42254 - The hearing aid shall include an inertial measurement unit. | UTICA-2760 - Verify by Inspection | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
|  |  | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 5 | passed
ElectroAcoustics | passed | MASTER-17316 - The hearing aid shall be capable of electrical operation using a battery from al... | UTICA-3249 - Zinc-air battery | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-2757 - Zinc-air battery | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-17383 - The hearing aid shall operate at all battery voltages from 1.1 to 1.5 VDC, with... | UTICA-3249 - Zinc-air battery | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
|  |  | UTICA-2757 - Zinc-air battery | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-17509 - The CIC hearing aid shall be available only in an Omnidirectional Configuration. | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | UTICA-2467 - The hearing aid with the Omnidirectional Configuration shall have one omnidirect... | UTICA-4132 - Verify by Inspection | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-59956 - The hearing aid shall detect the magnetic field from a telephone when held to th... | UTICA-3247 - Magnetic Switch | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-2762 - Magnetic Switch | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | UTICA-2495 - OSPL90: Each hearing aid listed in TABLE1 - PEAK OSPL90 AND PEAK GAIN PER MATRIX... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19750 - ACOUSTIC GAIN: Each hearing aid listed in TABLE1 - PEAK OSPL90 AND PEAK GAIN PER... |  |  | Executed | 2 | passed
ElectroAcoustics | passed | UTICA-2474 - FREQUENCY RANGE: Each hearing aid shall have a minimum frequency range as define... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-19808 - The hearing aid in-situ, occluded, frequency response shall be as flat as possib... | UTICA-2756 - Omni | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 4 | passed
|  |  | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-26435 - The hearing aid’s output spectrum measured at Full On Gain shall not exhibit pea... | UTICA-2759 - Sound Quality- Stability | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-4124 - Sound Quality- Stability | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-19812 - PEAK TO TROUGH (Nordic): The resonance is measured as the vertical distance betw... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 5 | passed
ElectroAcoustics | passed | MASTER-19813 - PEAK TO TROUGH (Australian): With all tone controls set for maximum gain, the pe... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 6 | passed
ElectroAcoustics | passed | MASTER-19817 - DISTORTION (Nordic): The hearing aid distortion shall comply with the following... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 7 | passed
ElectroAcoustics | passed | MASTER-19818 - DISTORTION (Australian): Total Harmonic Distortion shall not exceed the values s... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 8 | passed
ElectroAcoustics | passed | MASTER-19819 - IMD (Nordic): The hearing aid's intermodulation distortion shall comply with the... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 9 | passed
ElectroAcoustics | passed | MASTER-19820 - EIN - Omnidirectional (Nordic): The hearing aid shall meet or exceed the followi... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 10 | passed
ElectroAcoustics | passed | MASTER-19821 - EIN - Omnidirectional (Australian): The hearing aid shall meet or exceed the fol... | UTICA-3246 - Omni | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 11 | passed
ElectroAcoustics | passed | MASTER-19844 - POLARITY: The hearing aid shall preserve the polarity of the audio signal betwee... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4133 - Verify by Analysis | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-48451 - The hearing aid shall produce no audible artifacts that are deemed objectionable... | UTICA-4122 - Sound Quality- General / Audio Artifacts | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-2761 - Sound Quality- General / Audio Artifacts | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | UTICA-2476 - If the requirement for no audible artifacts while programming can't be achieved,... | UTICA-4122 - Sound Quality- General / Audio Artifacts | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | UTICA-19408 - The hearing aid shall have a minimum lifespan of 3 years under normal use. | UTICA-3249 - Zinc-air battery | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | UTICA-2519 - The materials used in the hearing aid shall not exceed 0.1 percent by weight in... | UTICA-4133 - Verify by Analysis | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-20085 - The materials used in the hearing aid shall not exceed 0.01 percent by weight in... | UTICA-2763 - Verify by Analysis | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 7 | passed
|  |  | UTICA-4133 - Verify by Analysis | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 3 | passed
ElectroAcoustics | passed | MASTER-20092 - The hearing aid shall produce no RF emissions greater than the following: * 30 d... | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-20099 - The hearing aid shall produce no RF emissions greater than the following: * 30 d... | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-20102 - The hearing aid shall meet or exceed Category M2 and T2 (when configured with a... | MASTER-45261 - EMC- RF Immunity | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-4128 - EMC- RF Immunity | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
ElectroAcoustics | passed | MASTER-20104 - The hearing aid shall meet or exceed the following: * User Compatibility for Omn... | MASTER-45261 - EMC- RF Immunity | EA_BTE_01 - Utica Power Plus BTE 13_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4128 - EMC- RF Immunity | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
ElectroAcoustics | passed | MASTER-20106 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 8 | passed
ElectroAcoustics | passed | MASTER-20113 - The hearing aid shall maintain wireless communication when exposed to an RF fiel... | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 4 | passed
ElectroAcoustics | passed | MASTER-20116 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 5 | passed
ElectroAcoustics | passed | MASTER-20120 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-4129 - EMC- ESD | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 1 | passed
|  |  | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 6 | passed
ElectroAcoustics | passed | MASTER-20126 - The hearing aid shall operate as specified herein, with no degradation in perfor... | UTICA-4129 - EMC- ESD | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
|  |  | UTICA-4130 - EMC- External Testing | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 7 | passed
ElectroAcoustics | passed | MASTER-20298 - The hearing aid shall pass the Peep Test as defined in Starkey Manufacturing Sui... | UTICA-4124 - Sound Quality- Stability | EA_CIC_01 - Utica 2.4 CIC_EA DVT1 | Executed | 2 | passed
Mechanical | passed | MASTER-17121 - The hearing aid faceplate and shell shall be available in the molded colors defi... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-23011 - The gap between the battery door and the faceplate, with and without a battery i... | UTICA-16715 - Quality Testing | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-17284 - The Right/Left Indicator shall be permanently attached to the hearing aid. | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 12 | passed
Mechanical | passed | MASTER-17285 - The Right/Left Indicator shall be no smaller than 1mm in diameter if round or 1m... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-19087 - Right/Left Indicator: CAD Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | UTICA-2459 - The hearing aid shall be compatible with zinc air batteries. | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 3 | passed
Mechanical | passed | MASTER-17308 - The hearing aid shall be compatible with zinc air batteries. | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 4 | passed
Mechanical | passed | MASTER-17313 - The hearing aid shall be capable of mechanically accommodating a battery from al... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-19461 - Battery: Size & Compatibility Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17389 - The battery door shall be labeled with a "+" sign (no quotes) that is no smaller... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 5 | passed
|  |  | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17410 - With the battery door in any position including the fully open position and the... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-19466 - Battery Door: Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 9 | passed
Mechanical | passed | MASTER-17412 - When the battery door is fully open, it shall position the battery entirely outs... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 4 | passed
|  |  | UTICA-19856 - Battery Door: Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17413 - When the battery door is fully open, it shall position the battery entirely outs... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 5 | passed
Mechanical | passed | MASTER-17420 - Attempting to close the battery door with the battery in the wrong orientation s... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 3 | passed
|  |  | UTICA-19856 - Battery Door: Analysis | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 3 | passed
Mechanical | passed | MASTER-17425 - With a battery installed, the battery door and its mechanism of attachment to th... | UTICA-16706 - Battery Door Cycling | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-22627 - Battery Door: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17428 - The hearing aid shall power on and off as specified herein when subjected to the... | UTICA-16716 - Durability Testing | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-22627 - Battery Door: Durability Testing | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-17430 - The battery door shall withstand a minimum of 10 removal and attachment cycles i... | UTICA-16716 - Durability Testing | Utica CIC ME DVT 1 | Executed | 2 | passed
Mechanical | passed | UTICA-2466 - The CIC hearing aid shall be available in the following control configurations:... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 6 | passed
Mechanical | passed | MASTER-17518 - The hearing aid design shall include protection to minimize foreign material ing... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 7 | passed
|  |  | UTICA-23260 - Microphone Protection: CAD Analysis/Physical Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-17519 - The microphone protection mechanism shall prevent the buildup of particulates an... | UTICA-16709 - Salt Fog | Utica CIC ME DVT 1 | Executed | 3 | passed
Mechanical | passed | MASTER-17657 - The hearing aid shall have the option of wax protection for the receiver. | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 8 | passed
Mechanical | passed | MASTER-17659 - The wax protection mechanism shall be compatible with RIC receiver cables and Cu... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 9 | passed
Mechanical | passed | MASTER-17660 - The wax protection mechanism shall be field replaceable. The use of a custom too... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 10 | passed
Mechanical | passed | MASTER-17662 - The hearing aid shall be compatible with the available HearClear products, part... | UTICA-16629 - Inspection | Utica CIC ME DVT 1 | Executed | 11 | passed
Mechanical | passed | MASTER-19880 - The hearing aid shall be fully operable after exposure to storage temperatures r... | UTICA-16712 - Storage | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23257 - Hi/Lo Storage Temp Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19906 - The hearing aid shall be fully operable during and after exposure to the light e... | UTICA-16711 - Light Sensitivity | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23237 - Light Sensitivity Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | UTICA-11615 - The hearing aid shall not have an operational failure from corrosion build-up on... | UTICA-16709 - Salt Fog | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-19907 - The hearing aid shall be fully operable after exposure to the heat and humidity... | UTICA-16710 - Damp Head Steady State | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23235 - DHSS/3yr & 5yr Lifespan Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19909 - The hearing aid exposure to the salt fog environment defined in Starkey specific... | UTICA-16709 - Salt Fog | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-23233 - Salt Fog Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-19916 - Plastic components shall meet the UV exposure test criteria defined in Starkey S... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23248 - UV Exposure Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19927 - The hearing aid parts shall meet the placement, alignment, color and quality cri... | UTICA-16708 - Internal Workmanship | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23224 - Final Visual Inspection | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19928 - The piece parts and assembly of the hearing aid shall meet the internal workmans... | UTICA-16708 - Internal Workmanship | Utica CIC ME DVT 1 | Executed | 2 | passed
|  |  | UTICA-23225 - Internal Workmanship | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-19930 - The color of the hearing aid faceplate and battery door shall be consistent when... | UTICA-16707 - Physical Inspection | Utica CIC ME DVT 1 | Executed | 6 | passed
Mechanical | passed | UTICA-19407 - The CIC hearing aid shall retain the removal handle with a break force of greate... | UTICA-16714 - RHT | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | UTICA-19408 - The hearing aid shall have a minimum lifespan of 3 years under normal use. | UTICA-22443 - Lifetime | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-20078 - All hearing aid parts shall be latex-free. | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-20079 - The hearing aid shall be free of the following phthalates: * DEHP * DBP * BBP *... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | UTICA-2519 - The materials used in the hearing aid shall not exceed 0.1 percent by weight in... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
Mechanical | passed | MASTER-20085 - The materials used in the hearing aid shall not exceed 0.01 percent by weight in... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23221 - European Directive 2011/65/EU (RoHS) - BTE13 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 2 | passed
Mechanical | passed | MASTER-20210 - All labeling on the hearing aid shall remain legible after exposure to the envir... | UTICA-22442 - Materials | Utica CIC ME DVT 1 | Executed | 1 | passed
|  |  | UTICA-23219 - Label Permanence & Abrasion Test | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 1 | passed
Mechanical | passed | MASTER-46245 - The packaging shall provide protection such that exposure to the shipping distri... | UTICA-22444 - Packaging | 0000 - Utica CIC Packaging Testing | Executed | 1 | passed
|  |  | UTICA-23194 - Shipping & Packaging (BTE13): ASTM D4169 | PPBTE - Utica 2.4GHz Pwr+ BTE13 ME DVT | Executed | 30 | passed
[5 empty] | Executed | 31 | passed
[5 empty] | Executed | 32 | passed
[5 empty] | Executed | 33 | passed
Quality | passed | MASTER-20082 - Any hearing aid part that comes into contact with the patient's skin during norm... | MASTER-39834 - Plastic Parts Evaluation Procedure | Utica Quality Test - Utica Pwr+ BTE | Executed | 5 | passed
Product Management | passed | MASTER-17275 - The hearing aid shall have a Right/Left Indicator. The professional and the pati... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 11 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 1 | passed
Product Management | passed | MASTER-17278 - The Right/Left Indicator shall be red to indicate Right, or blue to indicate Lef... |  |  | Executed | 2 | passed
Product Management | passed | MASTER-17288 - The Right/Left Indicator shall not be visible when the hearing aid is worn in th... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 14 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 3 | passed
Product Management | passed | MASTER-17301 - The hearing aid shall use a size 312 battery. |  |  | Executed | 4 | passed
Product Management | passed | MASTER-17388 - The battery door shall allow the patient to remove and replace the battery witho... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 15 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 5 | passed
Product Management | passed | MASTER-17393 - The battery door shall allow the patient to open and close the door without usin... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 16 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 6 | passed
Product Management | passed | MASTER-17394 - The battery door shall allow the patient to open and close the door without usin... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 17 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 7 | passed
Product Management | passed | MASTER-17411 - The battery door shall maintain its position, with and without a battery install... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 19 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 8 | passed
Product Management | passed | MASTER-17414 - The battery door or its attachment mechanism shall provide tactile feedback to i... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 20 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 9 | passed
Product Management | passed | MASTER-17415 - The battery door or its attachment mechanism shall provide a hard stop to indica... | UTICA-23584 - BTE SyRS Verification | Utica PP BTE Verification | Executed | 21 | passed
|  |  | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 10 | passed
Product Management | passed | UTICA-2462 - The battery door shall resist closing when the battery is installed upside down.... |  |  | Executed | 11 | passed
Product Management | passed | UTICA-2468 - The microphone cover shall be field replaceable. Note: The use of a custom tool... | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 12 | passed
Product Management | passed | MASTER-19926 - The hearing aid shall have no sharp edges or pointed protrusions that cut, poke,... | UTICA-23583 - CIC SyRS Verification | Utica 2_4 CIC Verification | Executed | 13 | passed
Regulatory | passed | MASTER-20092 - The hearing aid shall produce no RF emissions greater than the following: * 30 d... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 1 | passed
Regulatory | passed | MASTER-20099 - The hearing aid shall produce no RF emissions greater than the following: * 30 d... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 2 | passed
Regulatory | passed | MASTER-20106 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 3 | passed
Regulatory | passed | MASTER-20113 - The hearing aid shall maintain wireless communication when exposed to an RF fiel... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 4 | passed
Regulatory | passed | MASTER-20116 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 5 | passed
Regulatory | passed | MASTER-20120 - The hearing aid shall maintain its programmed settings and produce no output abo... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 6 | passed
Regulatory | passed | MASTER-20126 - The hearing aid shall operate as specified herein, with no degradation in perfor... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 7 | passed
Regulatory | passed | UTICA-2520 - The hearing aid shall comply with IEC 60601-2-66 (2019). | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 8 | passed
Regulatory | passed | MASTER-41126 - When in a single fault condition, the temperature of the exterior surface of the... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 9 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 7 | passed
Regulatory | passed | MASTER-20138 - The hearing aid shall meet the specific absorption rate (SAR) requirements for r... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 10 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 8 | passed
Regulatory | passed | MASTER-20143 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 11 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 9 | passed
Regulatory | passed | MASTER-20144 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 12 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 10 | passed
Regulatory | passed | MASTER-20145 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 13 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 11 | passed
Regulatory | passed | MASTER-20147 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 14 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 12 | passed
Regulatory | passed | MASTER-20148 - The hearing aid shall comply with US FCC Code of Federal Regulations, Title 47 T... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 15 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 13 | passed
Regulatory | passed | MASTER-20152 - The hearing aid shall comply with ERC/REC 70-03, Annex 3, frequency band A. | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 16 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 14 | passed
Regulatory | passed | MASTER-20153 - The hearing aid shall comply with Commission Decision 2009/381/EC for operation... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 17 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 15 | passed
Regulatory | passed | MASTER-20154 - The hearing aid shall comply with ETSI EN 300 328 for operation within the band... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 18 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 16 | passed
Regulatory | passed | MASTER-20158 - The hearing aid shall comply with the requirements of RSS-Gen that are applicabl... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 19 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 17 | passed
Regulatory | passed | MASTER-20159 - The hearing aid shall comply with RSS-247 for operation within the band of 2400-... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 20 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 18 | passed
Regulatory | passed | MASTER-20160 - The hearing aid shall comply with RSS-102 for operation within the band of 2400-... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 21 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 19 | passed
Regulatory | passed | MASTER-20161 - The hearing aid shall comply with RSP-100. | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 22 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 20 | passed
Regulatory | passed | MASTER-20164 - The hearing aid shall comply with the performance and methods of measurement of... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 23 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 21 | passed
Regulatory | passed | MASTER-20167 - The hearing aid shall comply with KN 301 489-1. Note: Testing is performed using... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 24 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 22 | passed
Regulatory | passed | MASTER-42289 - The hearing aid shall comply with KN 301 489-17. Note: Testing is performed usin... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 25 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 23 | passed
Regulatory | passed | MASTER-42290 - The hearing aid shall comply with Korean standard KN 300 328 for operation withi... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 26 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 24 | passed
Regulatory | passed | MASTER-20170 - The hearing aid shall comply with China Ministry of Industry and Information Tec... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 27 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 25 | passed
Regulatory | passed | MASTER-20172 - The hearing aid shall comply with Japan Association of Radio Industries and Busi... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 28 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 26 | passed
Regulatory | passed | MASTER-20174 - The hearing aid shall be qualified to the standards defined by the Bluetooth Spe... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 29 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 27 | passed
Regulatory | passed | MASTER-50185 - The hearing aid shall be labeled as specified in Starkey specification DOCS-19-1... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 30 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 28 | passed
Regulatory | passed | MASTER-20222 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 31 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 29 | passed
Regulatory | passed | MASTER-20223 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 32 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 30 | passed
Regulatory | passed | MASTER-20224 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 33 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 31 | passed
Regulatory | passed | MASTER-20225 - The Tinnitus Therapy feature labeling for the patient shall include appropriate... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 34 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 32 | passed
Regulatory | passed | MASTER-20226 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 35 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 33 | passed
Regulatory | passed | MASTER-20227 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 36 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 34 | passed
Regulatory | passed | MASTER-20228 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 37 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 35 | passed
Regulatory | passed | MASTER-20229 - The Tinnitus Therapy feature labeling for the patient shall include information... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 38 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 36 | passed
Regulatory | passed | MASTER-20230 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 39 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 37 | passed
Regulatory | passed | MASTER-20231 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 40 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 38 | passed
Regulatory | passed | MASTER-20232 - The Tinnitus Therapy feature labeling for the hearing professional shall include... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 41 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 39 | passed
Regulatory | passed | MASTER-20233 - The Tinnitus Therapy feature labeling shall include information on possible misu... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 42 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 40 | passed
Regulatory | passed | MASTER-20234 - The Tinnitus Therapy feature labeling shall include information on basic operati... | UTICA-24658 - Utica CIC Test Cases | Utica - Regulatory | Executed | 43 | passed
|  |  | UTICA-24660 - Utica PPBTE Test Cases | Utica - Regulatory | Executed | 41 | passed
Design For Manufacturing | passed | MASTER-48453 - The hearing aid circuit shall provide programming pads that are accessible prior... | UTICA-24619 - Programming Pad Inspection | Utica CIC test case - Programming Pad Inspection | Executed | N/A | passed